Validation Master Plan

Scope: Qualification and validation of a pre-vacuum steam sterilizer for terminal sterilization of medical device components and assemblies at the [Enter City, State here] facility. Includes IQ, OQ, PQ, requalification, and ongoing monitoring for all sterilization cycles defined in the URS.

Objective: Demonstrate the sterilizer is installed, operates, and performs reproducibly to meet URS requirements, regulatory expectations, and device sterility assurance targets.

References: AAMI ST79; ISO 17665; ISO 11138; FDA QSMR; FDA 21 CFR Part 11;[Enter Company name here] quality system procedures.

Roles and Responsibilities

- Validation Lead: coordinate validation activities, author protocols, compile final report.

- QA Manager: approve protocols, review results, final release decision.

- Engineering: installation, utilities, OQ execution support.

- Facilities: steam and utilities verification.

- Microbiology: BI/CI selection, incubation, interpretation.

- Production: PQ execution, operator training.

- IT: software configuration, 21 CFR Part 11 controls.

- Supplier Representative: FAT/SAT support and technical clarifications.

Qualification Strategy

- Execute IQ then OQ then PQ in sequence.

- Use supplier FAT/SAT data as supporting evidence.

- Apply risk-based approach to define mapping points, worst case loads, and requalification scope.

- Minimum PQ is three consecutive successful production loads.

Requalification Plan

- Triggers: relocation, major repair, unexplained BI failure, process change, CAPA.

- Frequency: annual full requalification unless risk assessment supports reduced scope.

- Scope: annual requalification includes selected OQ tests and one PQ load with BIs; post move requires full IQ OQ PQ unless risk assessment justifies reduced scope.

Deliverables: VMP, URS, IQ protocol and report, OQ protocol and report, PQ protocol and report, final validation report, traceability matrix, appendices.

User Requirement Specification URS (Map to PQ Protocol)

| ID | Requirement | Rationale | Acceptance Criterion | Test Method | Critical to Quality | Responsible |
| --- | --- | --- | --- | --- | --- | --- |
| URS-001 | Terminal sterilization of wrapped instruments, trays, porous loads, rigid containers, and lumened devices | Define scope of validated load types to ensure sterility assurance | PQ demonstrates sterility for each listed load type; BI negative for 3 consecutive PQ loads per load type | PQ with representative loads and BIs/CIs at worst-case locations | Y | QA; Microbiology; Production |
| URS-002 | Support for packaging formats including single wrap, double wrap, rigid containers, and peel pouches | Ensure validated cycles are compatible with production packaging | Packaging integrity maintained; no wet packs; residual moisture within limit after drying | PQ packaging inspection; residual moisture gravimetric test; CI evaluation | Y | Packaging Engineering; QA; Production |
| URS-003 | Maximum validated gross load mass of 50 kg and defined tray/cart configurations | Ensure chamber and steam penetration validated at production load limits | PQ loads up to 50 kg pass BI/CI and mapping shows acceptable temperatures | PQ with maximum mass loads and worst-case placement; temperature mapping in load | Y | Engineering; Validation Lead |
| URS-004 | Chamber dimensions and rack compatibility per DQ drawings | Confirm mechanical fit and validated load patterns | Physical dimensions match DQ drawings; racks fit and secure | IQ physical inspection and dimensional verification | N | Engineering |
| URS-005 | Support pre-vacuum cycles at 134°C/3 min and 132°C/4 min | Provide required lethality for different product families | OQ/PQ demonstrate required F₀ or BI pass for each cycle | OQ mapping and PQ BI challenge for each cycle | Y | Validation Lead; Microbiology |
| URS-006 | Support Bowie-Dick and dynamic air removal test cycles | Verify effective air removal and steam penetration for pre-vacuum operation | Bowie-Dick passes per standard; dynamic air removal within supplier limits | OQ Bowie-Dick tests and documented results | Y | Engineering; Validation Lead |
| URS-007 | User programmable cycles allowed only within defined limits and with role-based authorization | Prevent unauthorized or unsafe cycle changes | System enforces limits; audit trail records changes; only authorized roles can modify cycles | OQ software role tests; controlled edit attempts; audit trail verification | Y | IT; QA |
| URS-008 | Controller setpoint accuracy within ±0.5°C | Accurate control required to achieve validated lethality | Controller vs reference probe error ≤ ±0.5°C at steady state | OQ setpoint accuracy checks using calibrated reference thermocouple | Y | Engineering; Validation Lead |
| URS-009 | Chamber temperature uniformity within ±2.0°C across mapped points during exposure plateau | Uniform temperature ensures consistent lethality across load | All mapped points within ±2.0°C of setpoint in empty and loaded mapping | OQ temperature mapping with calibrated data loggers; 3 repeat runs | Y | Validation Lead |
| URS-010 | Temperature sensor redundancy for critical control points | Provide failover and detect sensor drift/failure | Redundant sensors installed; system alarms on sensor discrepancy | IQ verify sensor installation; OQ simulate discrepancy and verify alarm | Y | Engineering; QA |
| URS-011 | Pressure and vacuum sensor accuracy within ±2% full scale | Accurate pressure/vacuum measurement required for control and leak detection | Controller readings vs calibrated gauges within ±2% FS | OQ comparison to calibrated pressure/vacuum gauges during cycle | Y | Engineering |
| URS-012 | Vacuum capability to achieve ≤50 mbar and leak rate ≤5 mbar/min | Adequate vacuum required for effective air removal | Achieve ≤50 mbar and leak rate ≤5 mbar/min during OQ tests | OQ vacuum hold and leak rate tests with calibrated vacuum gauge | Y | Engineering |
| URS-013 | Steam quality with dryness fraction ≥0.95 and NCG ≤3.5% v/v | Steam quality affects heat transfer and lethality | Steam sampling meets dryness and NCG limits during OQ | Steam sampling and laboratory analysis per standard | Y | Facilities; QA |
| URS-014 | Steam inlet pressure range 3–6 bar and stable supply during cycles | Ensure consistent steam supply for validated cycles | Steam pressure within 3–6 bar during idle and cycle conditions | IQ steam inlet verification and OQ monitoring during cycles | Y | Facilities |
| URS-015 | Data logging of temperature, pressure, vacuum, cycle events, and alarms with retention ≥10 years | Maintain traceable records for compliance and investigations | Continuous recording with secure storage and backups; retention policy implemented | IQ verify backup config; OQ export and integrity checks; audit of retention | Y | IT; QA |
| URS-016 | Audit trail capturing user ID, timestamp, action, and reason for changes | Ensure accountability and support Part 11 where applicable | Audit trail entries present and immutable for all critical actions | OQ perform role actions and verify audit trail entries and immutability | Y | IT; QA |
| URS-017 | Role-based access control with unique user IDs and password policy | Prevent unauthorized access and changes | System enforces roles, password complexity, and lockout policies | OQ user role tests and attempted unauthorized actions | Y | IT |
| URS-018 | Electronic signature capability with traceable linkage to records where required | Support electronic approvals and regulatory requirements | Electronic signatures recorded with user ID and timestamp and linked to records | OQ electronic signature test and record verification | N | IT; QA |
| URS-019 | Data export in secure formats (CSV, PDF) with checksum or digital signature | Allow secure transfer of cycle data for review and archiving | Exported files readable; checksum matches; protected from tampering | OQ export test and checksum verification | N | IT |
| URS-020 | Alarm management with priority levels, logging, and operator guidance | Ensure timely response to faults and maintain safe operation | Alarms trigger notifications, are logged, and SOPs define responses | OQ simulate alarm conditions and verify notifications and logs | Y | Engineering; QA |
| URS-021 | Door interlock prevents opening under unsafe conditions and logs attempts | Prevent operator exposure and maintain sterile environment | Door cannot open until safe conditions met; attempts logged | IQ wiring check; OQ simulate unsafe condition and verify interlock | Y | Engineering; Safety |
| URS-022 | Emergency stop and fail-safe behavior on power loss | Ensure safe shutdown and prevent hazardous conditions | Emergency stop halts cycle and places system in safe state; power loss preserves data | OQ emergency stop test and simulated power loss test | Y | Engineering; QA |
| URS-023 | Controller firmware and software version control with supplier V&V evidence | Ensure validated software baseline and traceability of changes | Supplier provides V&V documentation and version matches DQ | IQ review supplier V&V and software version printout | Y | Procurement; IT; QA |
| URS-024 | Software validation and Part 11 compliance documentation where applicable | Demonstrate software meets regulatory requirements for electronic records | Software V&V report, Part 11 assessment, and test evidence present | OQ software function tests and Part 11 checklist | N | IT; QA |
| URS-025 | Continuous cycle reproducibility with statistical control | Demonstrate consistent performance over time | Key parameters show acceptable SD and capability; 3 consecutive PQ loads pass | OQ/PQ repeated cycles and statistical analysis | Y | Validation Lead; QA |
| URS-026 | BI type and placement strategy defined for each load type and cycle | Ensure appropriate biological challenge and worst-case coverage | BI species and D-value appropriate; placement covers worst-case; negative PQ results | PQ BI placement per plan and microbiology incubation with controls | Y | Microbiology; Validation Lead |
| URS-027 | Chemical indicator use and acceptance criteria for each pack and critical locations | Provide immediate visual confirmation of steam exposure | CIs show expected change for each pack and critical location in PQ | Visual inspection of CIs after each PQ load | Y | Production; QA |
| URS-028 | Residual moisture limit for packaged items after drying ≤3% by weight or CI dry-pack pass | Prevent wet packs that can compromise sterility during storage | Residual moisture ≤3% or CI indicates dry pack | Gravimetric moisture measurement or CI dry-pack evaluation | Y | QA; Production |
| URS-029 | Drying cycle performance and reproducibility | Ensure consistent drying across loads and pack types | Drying time achieves residual moisture acceptance across repeated runs | PQ drying tests with moisture measurement and CI checks | Y | Validation Lead; Production |
| URS-030 | Load configuration and worst-case placement definitions including lumens and dense packs | Validate penetration and heat transfer for realistic production patterns | Defined load patterns validated in PQ with BI negative at worst-case positions | PQ runs using defined load patterns and BI placement | Y | Production; Validation Lead |
| URS-031 | Minimum PQ requirement of three consecutive successful production loads per validated cycle | Provide statistical confidence in routine performance | Three consecutive PQ loads pass BI/CI and physical parameters | Execute three PQ loads under production conditions and document results | Y | Validation Lead; QA |
| URS-032 | Requalification triggers and schedule including annual and post-move requalification | Maintain validated state over equipment life and after changes | Requalification performed per VMP schedule and after triggers; records updated | Execute requalification protocol per VMP and document results | Y | QA; Validation Lead |
| URS-033 | Spare parts list and critical spares availability for validated components | Minimize downtime and ensure validated configuration can be maintained | Critical spares on site or on order prior to IQ completion | IQ documentation review and inventory check | Y | Procurement; Engineering |
| URS-034 | Preventive maintenance schedule and procedures that preserve validated state | Ensure maintenance does not invalidate validation or compromise performance | Approved PM schedule and procedures in place; maintenance records retained | Review PM SOPs and verify scheduled tasks in CMMS | Y | Maintenance; QA |
| URS-035 | Calibration intervals and traceability for all validation instruments and critical sensors | Ensure measurement accuracy and traceability to national standards | Calibration certificates traceable to NIST or equivalent and within interval | IQ collect calibration certificates; include in appendices | Y | Validation Lead; QA |
| URS-036 | Instrumentation mounting and data logger mounting method standardized and reproducible | Ensure mapping results are comparable and repeatable | Mounting fixtures documented and used consistently across mapping runs | OQ mapping documentation and photographic evidence of mounting | Y | Validation Lead; Engineering |
| URS-037 | Alarm logging and operator notification for utility failures such as steam loss or low pressure | Prompt detection and response to utility excursions that could affect cycles | Alarms logged and operator notified per SOP; automatic safe shutdown if required | OQ simulate utility failure and verify alarm and notification | Y | Facilities; Engineering; QA |
| URS-038 | Network security and cybersecurity controls for connected sterilizer controllers | Protect data integrity and prevent unauthorized access or tampering | Network segmentation, authentication, and encryption per IT policy; vulnerability assessment completed | IT security assessment and penetration test evidence | N | IT; Security |
| URS-039 | Backup and restore procedures for cycle data and system configuration | Ensure recoverability of validated configuration and records | Successful backup and restore test with verified data integrity | OQ backup/restore test and checksum verification | Y | IT; QA |
| URS-040 | Electronic and hardcopy record retention and archive indexing for validation records | Facilitate retrieval for audits and investigations | Records archived with index and retrieval procedure; retention meets policy (≥10 years) | Review archive and retrieval test | Y | QA; Document Control |
| URS-041 | Operator training program and competency assessment for validated operation and BI/CI interpretation | Ensure operators execute validated procedures correctly | Training records and competency assessments completed prior to PQ execution | Review training matrix and competency test results | Y | Training; Production; QA |
| URS-042 | SOPs for operation, cleaning, maintenance, BI/CI handling, and requalification | Standardize procedures to maintain validated state | Approved SOPs available and used; version controlled | IQ SOP review and evidence of use during PQ | Y | QA; Production |
| URS-043 | BI and CI lot control including lot certificates, D-values, and expiration tracking | Ensure biological challenge is valid and traceable | BI/CI lot certificates on file and within expiration; D-value documented | Review BI/CI lot certificates and attach to PQ report | Y | Microbiology; QA |
| URS-044 | Positive and negative controls for BI incubation included in each PQ run | Validate incubation conditions and BI viability | Positive control shows growth; negative control shows no growth | Incubation of control BIs per manufacturer instructions | Y | Microbiology |
| URS-045 | Traceability of each sterilized load to production batch and device lot numbers | Enable product release decisions and investigations | Load records include device lot numbers, operator, cycle ID, and BI/CI results | Review load records and traceability logs during PQ | Y | Production; QA |
| URS-046 | Capability to perform FAT and SAT with supplier and include FAT data in DQ package | Use supplier testing to reduce on-site qualification scope and verify design | FAT report provided and reviewed; SAT performed and accepted prior to IQ | Review FAT/SAT reports and attach to DQ/IQ appendices | Y | Procurement; Validation Lead; QA |
| URS-047 | Documentation of materials of construction and surface finish for chamber and load contact surfaces | Ensure compatibility with cleaning, sterilization, and regulatory requirements | Material certificates and surface finish specs provided and acceptable | IQ document review of material certificates and drawings | Y | Procurement; Engineering; QA |
| URS-048 | Cleaning and decontamination procedures for chamber and accessories | Prevent bioburden carryover and maintain validated state | Approved cleaning SOPs and evidence of cleaning during PQ and routine operation | Review cleaning logs and SOP adherence during PQ | Y | Production; QA |
| URS-049 | Environmental operating range for ambient temperature and humidity | Ensure sterilizer performance within facility environmental conditions | Sterilizer operates within specified ambient ranges without performance degradation | IQ verify ambient conditions and OQ monitor cycles under typical conditions | N | Facilities; Engineering |
| URS-050 | Floor loading, anchoring, and clearance requirements documented and met | Ensure safe installation and mechanical stability | Installation meets vendor floor loading and anchoring specs | IQ physical inspection and structural verification | N | Facilities; Engineering |
| URS-051 | Condensate handling and drainage capacity adequate for validated cycles | Prevent condensate backflow and ensure proper drainage during cycles | Drainage flows freely; condensate traps installed and functioning | IQ inspect drain routing and functional test during cycle | Y | Facilities |
| URS-052 | Audible and visual alarms for critical conditions with documented escalation path | Ensure operator awareness and timely response to critical events | Alarms audible/visible in operator area and escalation SOP defined | OQ alarm simulation and verification of notification path | Y | Engineering; Production; QA |
| URS-053 | Capability to perform partial requalification tests after minor repairs per VMP | Reduce downtime while ensuring validated state after limited changes | Defined partial requalification scope and tests executed with passing results | Execute partial requalification protocol and document results | Y | Validation Lead; QA |
| URS-054 | Capability to capture and store raw data files from controller and data loggers in secure archive | Preserve evidence for audits and investigations | Raw data files archived with checksum and access controls | OQ export and archive test with checksum verification | Y | IT; Validation Lead |
| URS-055 | Measurement uncertainty documented for all calibrated instruments used in validation | Ensure interpretation of mapping and test results accounts for instrument uncertainty | Calibration certificates include measurement uncertainty and are attached to reports | Review calibration certificates and include uncertainty in data analysis | Y | Validation Lead; QA |
| URS-056 | Capability to perform worst-case testing for lumened devices including longest lumen and smallest ID | Lumens present high risk for inadequate penetration and must be challenged | BI negative at lumen tips for validated lumen dimensions | PQ with lumened devices placed in worst-case positions and BI at lumen tip | Y | Microbiology; Production |
| URS-057 | Defined acceptance criteria and test methods for Bowie-Dick equivalent tests for larger chambers | Ensure air removal verification scales to chamber size and configuration | Bowie-Dick or equivalent test passes per defined criteria for chamber size | OQ run of BD or equivalent pack and documented pass criteria | Y | Validation Lead; Engineering |
| URS-058 | Documentation of change control process and criteria that trigger requalification | Maintain validated state and ensure appropriate response to changes | Change control procedure in place and applied; requalification triggers documented | Review change control records and VMP | Y | QA; Validation Lead |
| URS-059 | Defined CAPA process for deviations discovered during IQ/OQ/PQ and routine operation | Ensure systematic resolution and prevention of recurrence | CAPA initiated for significant deviations and closed with verification evidence | Review deviation and CAPA records associated with validation | Y | QA |
| URS-060 | Final validation report containing traceability matrix, raw data, BI/CI results, deviations, CAPA, and approvals | Provide complete evidence package for regulatory inspection and internal release | Final report complete, signed by Validation Lead, QA, Engineering, Microbiology, and Production | Compile and review final validation report and sign-off | Y | Validation Lead; QA |
| URS-061 | Material compatibility verification for each polymer, elastomer, adhesive, and coating used in devices | Some materials degrade under steam; validation must confirm no adverse change | No physical, dimensional, mechanical, or chemical degradation beyond defined limits after validated cycle(s) | Pre/post cycle inspection; tensile/elongation tests; dimensional checks; FTIR/extractables as applicable | Y | R&D; Materials Engineering; QA |
| URS-062 | Packaging material matrix validation for each packaging material and seal type | Different packaging behaves differently under steam and drying | Seal integrity maintained; CI/BI pass; no seal delamination or package failure across validated cycles | PQ with each packaging type; seal peel tests; CI/BI placement; visual inspection | Y | Packaging Engineering; QA; Production |
| URS-063 | Lumen and complex geometry matrix specifying max lumen length, min ID, and nested configurations | Lumens and complex geometries are worst-case for penetration | BI negative at lumen tips and in nested configurations for validated dimensions | PQ with lumened devices at worst-case positions and BI at lumen tips; include multiple lumen sizes | Y | Microbiology; Validation Lead |
| URS-064 | Validation for assembled and disassembled device states if both are processed | Assembly state can affect steam penetration and drying | Both states meet BI/CI and physical acceptance criteria when processed under validated cycles | PQ runs for each assembly state with BI/CI placement and post-cycle inspection | Y | Production; Validation Lead |
| URS-065 | Post-sterilization stability and shelf life testing requirements for sterilized product | Sterility must be maintained through intended shelf life and storage conditions | Stability study demonstrates sterility and product integrity for claimed shelf life under specified storage conditions | Accelerated and real-time stability testing; periodic sterility checks; packaging integrity over time | Y | Regulatory Affairs; QA; R&D |
| URS-066 | Feedwater and condensate quality limits and monitoring requirements | Water quality affects steam purity and equipment corrosion | Water quality parameters within defined limits and monitored per schedule | Water sampling and laboratory analysis; review facilities monitoring logs | Y | Facilities; QA |
| URS-067 | Altitude and ambient correction factors or cycle adjustments for facility elevation | Atmospheric pressure affects sterilization parameters and vacuum performance | DQ documents altitude adjustments; OQ demonstrates cycles meet lethality with corrections applied | DQ calculation and OQ verification runs at site altitude conditions | Y | Validation Lead; Engineering |
| URS-068 | HVAC and cleanroom interface requirements for sterilized product handling and storage | Post-sterilization handling environment can affect sterility and packaging | Storage and handling areas meet environmental class and monitoring requirements; records maintained | Environmental monitoring; HVAC performance verification; area qualification | Y | Facilities; QA |
| URS-069 | Exhaust, condensate disposal, and environmental compliance monitoring | Regulatory compliance for effluent and emissions | Condensate and exhaust meet local environmental regulations and monitoring shows compliance | Sampling and analysis of condensate/exhaust; review permits and monitoring logs | N | Facilities; EHS; QA |
| URS-070 | Firmware and software patch management policy including predeployment testing | Patches can change validated behavior; controlled deployment required | Patches tested in controlled environment; change control completed; revalidation scope assessed before deployment | Patch test report; change control record; post-patch verification tests | N | IT; Validation Lead; QA |
| URS-071 | Time synchronization using NTP or equivalent for audit trail integrity | Accurate timestamps essential for traceability and audits | Controller and archive systems synchronized to approved time source; drift within tolerance | IQ verify NTP config; periodic timestamp drift checks | N | IT; QA |
| URS-072 | Encryption and secure transport for data in transit and at rest | Protect data integrity and confidentiality | Data encrypted per IT policy; secure protocols used; access controls enforced | IT security assessment and verification of encryption settings | N | IT; Security |
| URS-073 | MES/LIMS or third-party integration requirements and data mapping | Ensure correct transfer of cycle and lot data to enterprise systems | Data fields mapped; transfers validated; reconciliation tests pass | Integration test scripts; sample transfers; reconciliation checks | N | IT; Production; QA |
| URS-074 | Disaster recovery objectives for controller and data systems including RTO/RPO | Ensure recoverability of validated configuration and records after outage | Backup and restore meet defined RTO/RPO; restore test successful | OQ backup/restore test and documented recovery time | N | IT; QA |
| URS-075 | Lot release rules and automated hold/release logic for loads | Define when product can be released to prevent premature release | Loads held until BI/CI and approvals complete; automated holds function as designed | Review release logic; simulated holds; release workflow tests | Y | QA; Production; IT |
| URS-076 | Statistical sampling plans for large loads or batch releases including acceptance numbers and actions on failures | Provide statistically justified sampling for release decisions | Sampling plan documented and statistically justified; implemented during PQ and routine release | Statistical plan document and execution during PQ and production | Y | QA; Biostatistics |
| URS-077 | Indicator selection criteria including BI species, CI classes, and lot acceptance rules | Ensure appropriate challenge and immediate verification | BI species and CI classes specified; lot certificates reviewed and accepted before use | Review BI/CI lot certificates and include in PQ appendices | Y | Microbiology; QA |
| URS-078 | Environmental monitoring requirements for sterilized product storage areas | Maintain storage conditions that preserve sterility and packaging integrity | Monitoring frequency and alert thresholds defined and met; corrective actions documented | Environmental monitoring logs and trend review | N | QA; Facilities |
| URS-079 | Labeling and electronic traceability fields required on load records and pack labels | Ensure product traceability for release and investigations | Load records include device lot numbers, operator, cycle ID, BI/CI results, and timestamp | Review load records and label samples during PQ | Y | Production; QA |
| URS-080 | Critical spare lead times and minimum on-hand spares for key components | Minimize downtime and maintain validated configuration | Critical spares list with lead times and minimum stock levels approved and available | IQ inventory check and procurement records | N | Procurement; Maintenance; Engineering |
| URS-081 | Maintenance actions that require partial or full requalification defined and documented | Some maintenance can affect validated state and must trigger requalification | List of maintenance actions and associated requalification scope documented and approved | Review maintenance SOPs and change control records | Y | Maintenance; QA; Validation Lead |
| URS-082 | Consumables qualification requirements for filters, gaskets, trays, and indicators | Consumables affect performance and must be controlled | Supplier qualification, certificates of conformity, and acceptance testing for consumables | Review supplier certificates and perform incoming inspection tests | Y | Procurement; QA |
| URS-083 | End-of-life and decommissioning criteria and records retention for retired equipment | Ensure proper disposition and record keeping for decommissioned validated equipment | Decommissioning procedure executed and records archived per retention policy | Review decommissioning records and archive index | N | QA; Facilities |
| URS-084 | Regulatory reporting triggers and timelines for events such as widespread BI failures | Meet regulatory obligations for reportable events | Events that trigger reporting defined and procedures in place to notify regulators within timelines | Review regulatory reporting SOP and sample event simulation | Y | Regulatory Affairs; QA |
| URS-085 | Emergency procedures for critical failures including product containment and notification | Protect patient safety and product integrity during incidents | Emergency SOPs available, personnel trained, and drills performed | Review SOPs, training records, and evidence of drills | Y | QA; Production; EHS |
| URS-086 | Audit evidence readiness checklist specifying minimum dataset and format for inspections | Ensure rapid response to audits and inspections | Audit pack template available and validated records organized per template | Mock audit of validation package and retrieval time test | N | QA; Document Control |
| URS-087 | Local and regional regulatory requirements specific to facility location documented and met | Compliance with state and local environmental, waste, and safety rules | Permits and local regulatory requirements identified and evidence of compliance available | Regulatory review and permit documentation check | N | Regulatory Affairs; EHS |
| URS-088 | Consumable and spare part change control including approved alternatives and requalification impact | Substituting parts can change performance and must be controlled | Approved list of alternatives and documented impact assessment for each substitution | Change control records and verification tests for substituted parts | Y | Procurement; QA; Engineering |
| URS-089 | Training and competency requalification frequency for operators and maintenance staff | Maintain operator competence over time and after changes | Training schedule and competency assessments completed at defined intervals and after major changes | Review training matrix and competency test records | Y | Training; QA |
| URS-090 | Trending and periodic review of BI/CI failures, deviations, and maintenance events to detect systemic issues | Early detection of trends prevents recurrence and maintains validated state | Periodic trend reports generated and reviewed; actions taken for adverse trends | Quarterly trend analysis report and management review minutes | Y | QA; Validation Lead |
| URS-091 | Validation of alternative sterilization cycles or emergency cycles and their requalification requirements | Emergency cycles may be needed but must be controlled and validated before routine use | Emergency cycle defined, risk assessed, and limited validation performed before use; requalification plan documented | Risk assessment; limited OQ/PQ tests; approval via change control | N | Validation Lead; QA |
| URS-092 | Requirements for third-party service providers performing maintenance or calibration on critical components | External service providers must maintain validated state and traceability | Service provider qualifications, certificates, and work records on file; work under approved procedures | Review service provider qualifications and sample service reports | N | Maintenance; Procurement; QA |
| URS-093 | Requirements for validation of sterilizer accessories such as racks, trays, and containers | Accessories affect load configuration and heat transfer | Accessories validated as part of load patterns and included in PQ | PQ runs including accessories and accessory compatibility tests | Y | Production; Validation Lead |
| URS-094 | Requirements for handling and disposition of nonconforming sterilized product | Ensure consistent containment and investigation of suspect product | Nonconforming product procedure in place and followed; records retained | Review nonconforming product logs and sample investigations | Y | QA; Production |
| URS-095 | Requirements for periodic verification of steam trap and separator performance | Steam quality depends on proper trap and separator function | Periodic checks show traps and separators functioning within vendor limits; records maintained | Facilities inspection and steam quality sampling after maintenance | Y | Facilities; Maintenance |
| URS-096 | Requirements for measurement system analysis for critical measurement instruments used in validation | Ensure measurement systems are capable and uncertainty understood | MSA performed for data loggers, thermocouples, and pressure gauges with acceptable results | Conduct MSA studies and include results in OQ report | Y | Validation Lead; QA |
| URS-097 | Requirements for periodic review of validation status and VMP updates | Keep validation strategy current with process and regulatory changes | Annual review of validation status and VMP with documented updates as needed | Management review and VMP revision record | N | QA; Validation Lead |
| URS-098 | Requirements for supplier quality agreements covering FAT, SAT, spare parts, and software support | Define supplier responsibilities and deliverables that affect validated state | Signed supplier quality agreement on file covering required items | Review supplier agreement and evidence of compliance | N | Procurement; QA |
| URS-099 | Requirements for validation of sterilizer performance under seasonal or extreme ambient conditions if applicable | Ambient extremes can affect performance and must be assessed | OQ demonstrates acceptable performance within defined ambient extremes or mitigations documented | OQ runs under simulated ambient extremes or engineering analysis with verification | N | Facilities; Validation Lead |
| URS-100 | Requirements for documentation of measurement uncertainty propagation into acceptance decisions | Ensure acceptance criteria account for instrument uncertainty | Data analysis includes measurement uncertainty and acceptance decisions justified accordingly | Include uncertainty calculations in OQ/PQ reports and decision rationale | Y | Validation Lead; QA |
| URS-101 | Pre-cleaning and bioburden limits for incoming loads | High pre-sterilization bioburden can compromise cycle lethality and BI interpretation | Defined maximum allowable bioburden per device type; pre-cleaning SOPs executed and records show compliance | Incoming sample bioburden testing; review pre-cleaning logs; periodic verification sampling | Y | Production; Microbiology; QA |
| URS-102 | Approved detergents and cleaning agents compatibility list | Residual cleaning agents can affect sterilization and packaging integrity | Only approved agents used; compatibility testing shows no adverse effects | Chemical compatibility tests; residue analysis; process audits | N | R&D; QA; Production |
| URS-103 | Validation of sterilizer performance with alternative steam sources (central plant vs local generator) | Different steam sources can vary in quality and pressure stability | OQ steam sampling and cycle verification pass for each steam source used | OQ runs and steam quality sampling for each source; compare to URS limits | Y | Facilities; Validation Lead |
| URS-104 | UPS and emergency power requirements for safe shutdown and data preservation | Power interruptions can corrupt cycle data and create unsafe conditions | UPS supports safe shutdown and data preservation for defined interval; successful simulated power loss test | OQ simulated power loss and UPS failover test with data integrity verification | N | IT; Facilities; Engineering |
| URS-105 | Electromagnetic compatibility and grounding requirements | EMI can affect controller operation and data integrity | Sterilizer meets applicable EMC standards and grounding requirements; no controller anomalies during EMI tests | Review supplier EMC documentation and perform site EMI/grounding checks if required | N | Engineering; IT |
| URS-106 | Noise and vibration limits for installation area | Excessive vibration or noise can affect equipment and operator safety | Measured noise and vibration within acceptable occupational and equipment limits | IQ environmental measurements and review of vendor specifications | N | Facilities; EHS |
| URS-107 | Validation of sterilizer performance across seasonal ambient extremes if facility experiences them | Ambient extremes can affect drying and control loops | OQ demonstrates acceptable performance within defined ambient extremes or mitigations documented | OQ runs under simulated ambient extremes or engineering analysis with verification | N | Facilities; Validation Lead |
| URS-108 | Measurement system analysis for critical instruments including repeatability and reproducibility | Ensure measurement systems are capable and uncertainty quantified | MSA results meet predefined capability criteria; uncertainty included in acceptance decisions | Conduct MSA studies (Gage R&R) and document results in OQ report | Y | Validation Lead; QA |
| URS-109 | Sensor drift detection and recalibration triggers | Sensor drift can invalidate mapping and cycle control | Drift thresholds defined; automatic or scheduled recalibration triggered when exceeded | OQ simulate drift scenarios; implement monitoring and calibration schedule | Y | Engineering; QA |
| URS-110 | Firmware rollback and recovery procedures | Ability to revert to validated firmware if an update causes issues | Documented rollback procedure tested successfully without data loss | IQ verify rollback capability; OQ perform controlled rollback test in non-production environment | N | IT; Supplier; QA |
| URS-111 | Validation of remote access controls and restrictions for vendor support | Remote access can introduce security and change risks | Remote access restricted, logged, time-limited, and approved via change control; test logs show access events | OQ simulate vendor remote session and verify logging and access controls | N | IT; QA; Supplier |
| URS-112 | Mobile device and tablet access policy for operator interfaces | Mobile access can affect data security and operator behavior | Mobile access controlled per IT policy; audit trail captures mobile user actions | OQ role and access tests using mobile devices; verify audit trail | N | IT; QA |
| URS-113 | Validation of load release automation and interlocks with MES/LIMS | Automated release must not bypass microbiology or QA checks | Automated release logic enforces holds until BI/CI and approvals complete; integration tests pass | Integration test scripts; simulated holds; reconciliation checks | Y | IT; QA; Production |
| URS-114 | Sample retention policy for sterilized product and test samples | Retained samples support investigations and stability testing | Retention locations, durations, and sample sizes defined and implemented | Review retention logs and physical verification of retained samples | N | QA; Production |
| URS-115 | Validation of sterilizer performance after major maintenance activities | Major maintenance can alter validated performance | Defined post-maintenance verification tests executed and passed before return to production | Post-maintenance verification protocol and test execution | Y | Maintenance; Validation Lead; QA |
| URS-116 | Obsolescence management and spare part lifecycle planning | Supplier part obsolescence can affect long-term maintainability | Obsolescence plan and approved alternatives documented; impact assessments performed | Review supplier lifecycle documentation and change control records | N | Procurement; Engineering; QA |
| URS-117 | Validation with alternative packaging orientations and stacking patterns | Orientation and stacking affect steam penetration and drying | PQ demonstrates acceptable BI/CI results for validated orientations and stacking patterns | PQ runs with alternative orientations and stacking patterns; BI placement at worst-case points | Y | Production; Validation Lead |
| URS-118 | Validation when processing mixed loads (different device types in same cycle) | Mixed loads can create complex worst-case scenarios | Defined mixed load patterns validated with BI/CI and meet acceptance | PQ mixed load runs with BI/CI placement and analysis | Y | Production; Validation Lead |
| URS-119 | Handling and validation requirements for implantable devices including additional documentation | Implants may have stricter release and documentation requirements | Additional implant handling SOPs and release criteria defined and met; traceability ensured | PQ with implant loads, enhanced documentation, and regulatory review | Y | Regulatory Affairs; QA; Production |
| URS-120 | Pyrogen awareness and controls for devices where endotoxin is a concern | Steam sterilization does not remove endotoxin; controls must prevent introduction | Processes prevent endotoxin introduction; endotoxin testing performed where required | Process controls review and endotoxin testing on representative lots if applicable | Y | QA; R&D; Microbiology |
| URS-121 | Validation after relocation within facility including transport vibration effects | Movement can affect alignment, seals, and sensors | Post-move IQ/OQ executed per VMP; tests pass before production use | Full or partial IQ/OQ per VMP and documented transport inspection | Y | Validation Lead; Engineering |
| URS-122 | Multi-site consistency requirements if same sterilizer model used at other facilities | Ensure consistent validation approach and comparability across sites | Site-to-site DQ/OQ comparability assessment and harmonized protocols where applicable | Cross-site DQ comparison; FAT/SAT review; site OQ alignment | N | Corporate Validation; QA |
| URS-123 | Periodic supplier audits for critical component and software suppliers | Supplier quality affects validated state and long-term support | Supplier audit schedule and completed audit reports on file for critical suppliers | Supplier audit reports and corrective action follow-up review | N | Procurement; QA |
| URS-124 | Validation of sterilizer accessories cleaning and sterilization compatibility | Accessories may retain moisture or bioburden and affect cycles | Accessories validated for cleaning and sterilization; no adverse effects observed | Accessory cleaning validation and PQ inclusion tests | Y | Production; Validation Lead |
| URS-125 | Documentation of raw data provenance and chain of custody for electronic records | Ensure authenticity and integrity of validation evidence | Provenance metadata captured for raw files and chain of custody documented | OQ export and metadata verification; audit of file handling procedures | Y | IT; QA |
| URS-126 | Anonymization or redaction rules for records shared externally | Protect confidential product or patient information while enabling review | Redaction rules defined and applied; redacted records retain necessary validation evidence | Review of redacted sample records and verification of retained evidence | N | QA; Regulatory Affairs; Legal |
| URS-127 | Validation with different water hardness levels | Water hardness affects steam generation and scale formation | Steam quality and cycle performance acceptable across defined water hardness range or mitigations in place | Steam sampling and performance verification with representative water quality or controls | N | Facilities; Validation Lead |
| URS-128 | Periodic verification of steam trap and separator maintenance effectiveness | Proper trap/separator function is critical to steam quality | Periodic checks show traps and separators functioning within vendor limits; records maintained | Facilities inspection and steam quality sampling after maintenance | Y | Facilities; Maintenance |
| URS-129 | Validation under controlled humidity variations during drying phase | Ambient and chamber humidity affect drying efficiency | Drying performance meets residual moisture limits across defined humidity range or mitigations documented | PQ drying tests under varied humidity conditions or engineering analysis with verification | N | Validation Lead; Facilities |
| URS-130 | Validation when processing loads with embedded electronics or sensors declared sterilizable | Electronics may be sensitive to steam and temperature cycles | Manufacturer declarations and compatibility testing show no functional degradation post-cycle | Functional testing of electronics pre/post sterilization and vendor compatibility evidence | N | R&D; QA; Supplier |
| URS-131 | Validation with different packaging sterilant indicators and interpretation criteria | Different CI classes provide different levels of information and must be interpreted correctly | CI class selection justified per load type and CI results interpreted per manufacturer guidance; documented in PQ | PQ with specified CI classes and documented interpretation | Y | Microbiology; QA |
| URS-132 | Validation when using alternative sterilization cycles for special materials | Some materials require non-standard cycles that must be validated | Alternative cycles validated with BI/CI and material compatibility testing; approved via change control | OQ/PQ for alternative cycles and material testing | Y | Validation Lead; R&D; QA |
| URS-133 | Periodic verification of controller clock and timestamp integrity including DST and leap seconds | Timestamp integrity is critical for audit trails and traceability | Controller timestamps correct and synchronized; documented checks performed periodically | IQ verify NTP sync; periodic timestamp verification logs | N | IT; QA |
| URS-134 | Validation with different sterilizer door seal materials and replacement parts | Seal material affects vacuum integrity and leak rates | Replacement seals validated for leak rate and cycle performance; records of part numbers and lot numbers maintained | OQ leak tests and PQ verification after seal replacement | Y | Maintenance; QA |
| URS-135 | Validation when using load carriers with different thermal mass (aluminum vs stainless) | Thermal mass affects heating/cooling profiles and may change lethality | PQ demonstrates acceptable temperatures and BI/CI results for each carrier type | OQ/PQ runs with each carrier type and mapping in load | Y | Production; Validation Lead |
| URS-136 | Validation when processing loads with adhesives, labels, or inks that may be affected by steam | Adhesives and inks can migrate or degrade, affecting product quality | No unacceptable migration, delamination, or functional degradation post-cycle | Pre/post cycle inspection; adhesion tests; chemical analysis if required | N | R&D; QA |
| URS-137 | Validation with nested packaging configurations (packs inside containers) | Nested configurations can impede steam penetration and drying | BI/CI negative at worst-case nested locations and drying acceptance met | PQ with nested configurations and BI placement at nested cores | Y | Production; Validation Lead |
| URS-138 | Validation when processing loads with variable initial moisture content | Initial moisture affects cycle energy balance and drying | PQ demonstrates acceptable performance across defined initial moisture range or pre-conditioning required | PQ runs with controlled initial moisture levels and moisture measurements post-cycle | N | Validation Lead; Production |
| URS-139 | Validation when processing loads with biological materials or tissues if applicable | Biological materials present unique contamination and safety considerations | Special handling SOPs, containment, and validation tests defined and passed; regulatory approvals obtained if required | Risk assessment; containment verification; specialized PQ with appropriate controls | Y | R&D; QA; EHS |
| URS-140 | Continuous improvement feedback loop from production to validation including lessons learned and protocol updates | Field experience should inform validation lifecycle and reduce recurrence of issues | Periodic review meetings, documented lessons learned, and VMP/URS updates where justified | Management review minutes; VMP revision records; evidence of implemented improvements | N | QA; Validation Lead; |

Functional Requirement Specification FRS (Map to OQ Protocol)

| ID | FRS (Requirement) | Rationale | Acceptance Criterion | Test Method | Critical to Quality | Responsible |
| --- | --- | --- | --- | --- | --- | --- |
| FRS-001 | Support validated sterilization workflows for wrapped instruments, trays, porous loads, rigid containers, and lumened devices. | Translate user scope into functional behaviors to execute and record validated sterilization for each load type. | Selectable workflows for each load type; execution records include cycle ID, load map, BI/CI placements, pass/fail flags. | Create and execute workflows for each load type; review generated records for required fields and BI/CI linkage. | Y | Validation Lead; IT; Production |
| FRS-002 | Process and dry single/double wrap, rigid containers, and peel pouches without damaging packaging. | Ensure functional controls to select appropriate cycle and drying parameters per packaging format. | Packaging-specific cycle options and drying profiles available; CI/pack integrity fields captured. | Execute packaging-specific cycles; verify CI response, seal integrity checks, and drying profile logs. | Y | Production; Validation Lead; QA |
| FRS-003 | Maintain cycle performance up to maximum gross load mass (50 kg). | Ensure cycle parameters remain effective at maximum production load. | System blocks start if load >50 kg; cycle logs show parameters within limits for 50 kg loads. | Attempt start with >50 kg (expect block); run 50 kg cycle and verify logged parameters and BI/CI outcomes. | Y | Engineering; Validation Lead |
| FRS-004 | Accept specified chamber dimensions and rack configurations and enforce load pattern templates. | Map physical fixtures to load templates to ensure repeatable placement and mapping. | Load templates match DQ drawings; system enforces template selection for given rack type. | Create/select load template; verify HMI prompts and record template used in cycle log. | N | Engineering; Production |
| FRS-005 | Execute and control pre-vacuum cycle recipes (134°C/3 min and 132°C/4 min) reproducibly. | Provide recipe management and execution to meet lethality targets. | Recipes stored/protected; execution logs include setpoints, plateau times, and F₀ calculations. | Run each recipe; verify setpoint attainment, plateau duration, and F₀ calculation in cycle report. | Y | Validation Lead; Engineering; IT |
| FRS-006 | Provide built-in Bowie-Dick and dynamic air removal diagnostic routines and reporting. | Enable routine verification of air removal performance for pre-vacuum operation. | BD routine available; produces pass/fail output stored in logs. | Execute BD routine per standard; verify pass/fail output and log entry. | Y | Engineering; Validation Lead |
| FRS-007 | Restrict user programmable cycle edits to authorized roles and enforce parameter limits. | Prevent unauthorized or unsafe cycle changes that could compromise sterility. | RBAC enforced for recipe edits; out-of-range parameters blocked; attempts logged. | Attempt recipe edit as unauthorized user (expect deny); edit as authorized user and verify audit entry. | Y | IT; QA |
| FRS-008 | Maintain controller setpoint control within ±0.5°C during exposure plateau. | Ensure control accuracy required to achieve validated lethality. | Plateau temperature within ±0.5°C of setpoint for sustained period. | Compare controller readout to calibrated reference probe during plateau; verify within ±0.5°C. | Y | Engineering; Validation Lead |
| FRS-009 | Achieve and report temperature uniformity within ±2.0°C across mapped points during plateau. | Ensure consistent sterilization across the entire load. | Mapping reports show all mapped points within ±2.0°C of setpoint for empty and loaded conditions. | Perform mapping runs and analyze logger data for uniformity compliance. | Y | Validation Lead; Engineering |
| FRS-010 | Monitor redundant temperature sensors and alarm on discrepancy; provide failover logic. | Provide redundancy detection to avoid single-sensor reliance and detect failures. | System detects sensor discrepancy beyond threshold, raises alarm, and follows defined fallback behavior. | Simulate sensor discrepancy and verify alarm, log entry, and fallback behavior. | Y | Engineering; QA |
| FRS-011 | Measure and log pressure and vacuum with accuracy within ±2% full scale. | Accurate pressure/vacuum measurement required for control and leak detection. | Logged pressure/vacuum values match calibrated reference within ±2% FS across operating range. | Compare controller readings to calibrated gauges at multiple points; verify within tolerance. | Y | Engineering |
| FRS-012 | Provide vacuum generation and hold functions to achieve ≤50 mbar and enable leak testing. | Ensure adequate vacuum for effective air removal and leak verification. | System achieves ≤50 mbar and supports automated leak test routine with calculated leak rate. | Execute vacuum and leak test routine; verify vacuum level and leak rate calculation. | Y | Engineering |
| FRS-013 | Support steam quality monitoring (dryness fraction and NCG) and sampling interfaces. | Steam quality affects heat transfer and lethality; monitoring required. | Sampling ports present; lab results linked to cycle records and meet dryness/NCG limits. | Collect steam sample via port, attach lab report to cycle record, and verify acceptance fields. | Y | Facilities; QA |
| FRS-014 | Monitor and log steam inlet pressure and enforce acceptable range (3–6 bar). | Ensure consistent steam supply for validated cycles. | Inlet pressure logged continuously; system alarms and blocks cycles outside 3–6 bar. | Simulate inlet pressure excursions and verify alarms and cycle interlocks. | Y | Facilities; Engineering |
| FRS-015 | Continuously log temperature, pressure, vacuum, cycle events, and alarms; support secure archival with ≥10-year retention. | Maintain traceable records for regulatory compliance and investigations. | Continuous logs generated, exported, archived with checksum, and retention policy enforced. | Run cycles, export logs, verify checksum, and confirm archive retention settings. | Y | IT; QA |
| FRS-016 | Provide immutable audit trail capturing user ID, timestamp, action, and reason for critical actions. | Ensure accountability and support Part 11 compliance where applicable. | Audit entries immutable, include required metadata, and are searchable. | Perform critical actions and verify audit entries and immutability. | Y | IT; QA |
| FRS-017 | Enforce role-based access control with unique user IDs and password policies. | Prevent unauthorized access and changes to validated system. | RBAC enforced; password complexity and lockout policies active; user provisioning logged. | Create test accounts and verify policy enforcement and logs. | Y | IT |
| FRS-018 | Support electronic signatures linked to records where required. | Enable electronic approvals and regulatory requirements. | E-signature captured with user ID, reason, timestamp and linked to the record. | Apply e-signature to test record and verify metadata and linkage. | N | IT; QA |
| FRS-019 | Export cycle data in secure formats (CSV, PDF) with integrity verification (checksum/digital signature). | Allow secure transfer of cycle data for review and archiving. | Exported files include checksum/digital signature and match archived originals. | Export sample cycle file, compute/verify checksum, and validate signature. | N | IT |
| FRS-020 | Classify, log, and present alarms with priority levels and operator guidance. | Ensure timely and appropriate operator response to faults. | Alarms categorized by priority, logged, and HMI displays operator actions and SOP links. | Trigger alarms of different priorities and verify HMI guidance, logs, and escalation behavior. | Y | Engineering; QA |
| FRS-021 | Enforce door interlock logic preventing opening under unsafe conditions and log attempts. | Prevent operator exposure and preserve cycle integrity. | Door remains locked until safe conditions; open attempts logged with user/timestamp. | Simulate unsafe condition and verify door locked and log entry. | Y | Engineering; Safety |
| FRS-022 | Implement emergency stop and safe-state logic with data preservation on power loss. | Protect personnel and preserve cycle data during outages. | E-stop triggers safe state; UPS preserves controller data; power-loss events logged. | Activate E-stop and simulate power loss; verify safe state and data integrity. | Y | Engineering; IT |
| FRS-023 | Record and enforce approved firmware/software versions and link to supplier V&V evidence. | Ensure only validated software runs on system. | Installed firmware version recorded and linked to supplier V&V package. | Verify firmware version and presence of supplier V&V documentation. | Y | Procurement; IT; QA |
| FRS-024 | Provide software validation artifacts and Part 11 functional controls where applicable. | Demonstrate software meets regulatory expectations for electronic records. | V&V deliverables present; Part 11 controls implemented and tested. | Execute software function tests per V&V and verify Part 11 controls. | N | IT; QA |
| FRS-025 | Provide SPC tools and reproducibility monitoring for key cycle parameters. | Detect drift and ensure consistent performance over time. | SPC charts available; system flags out-of-control conditions and computes SD/Cp/Cpk if required. | Run repeated cycles, generate SPC charts, and verify alerts for out-of-control. | Y | Validation Lead; QA |
| FRS-026 | Support BI placement plans and record BI locations per load. | Ensure worst-case coverage and traceability of biological challenge. | BI placement templates available; BI locations recorded in load record. | Create BI plan, execute placement, and verify BI location entries in cycle record. | Y | Microbiology; Production |
| FRS-027 | Capture CI placement and results per pack and critical locations. | Provide immediate visual verification of steam exposure. | CI placement recorded and CI results logged with pass/fail flags. | Place CIs, run cycle, and verify CI result entries and flags in system. | Y | Production; QA |
| FRS-028 | Provide drying cycle control and residual moisture measurement workflow. | Ensure packs are dry for storage and prevent wet packs. | Drying profiles logged; residual moisture results captured and compared to ≤3% acceptance. | Run drying cycle, perform gravimetric moisture test, and verify results logged and within acceptance. | Y | Validation Lead; Production; QA |
| FRS-029 | Ensure reproducible drying performance across repeated cycles and loads. | Prevent variability that could produce wet packs. | Repeated drying cycles produce residual moisture within acceptance and SD within defined limits. | Execute repeated drying cycles and analyze moisture results for reproducibility. | Y | Validation Lead; Production |
| FRS-030 | Enforce load configuration templates and worst-case placement recording (including lumens). | Ensure consistent worst-case testing and mapping traceability. | System requires selection of load template and records placement coordinates for each cycle. | Select load template, run cycle, and verify placement coordinates and template ID in cycle record. | Y | Production; Validation Lead |
| FRS-031 | Provide PQ workflow enforcing three consecutive successful production loads per validated cycle. | Ensure statistical confidence for routine performance. | PQ workflow enforces three-load sequence, captures BI/CI results, and prevents release until all pass. | Execute PQ workflow and verify enforcement of three-load sequence and release gating. | Y | Validation Lead; QA |
| FRS-032 | Integrate requalification scheduling and trigger logic with change control. | Ensure timely requalification after defined triggers. | Requalification schedule present; change control triggers requalification workflow automatically. | Create change control event and verify requalification trigger and scheduling. | Y | QA; Validation Lead |
| FRS-033 | Maintain spare parts master data and minimum stock alerts. | Reduce downtime and preserve validated configuration. | Spare parts list with min stock levels and automated reorder alerts. | Simulate part consumption and verify reorder alert generation. | Y | Procurement; Maintenance |
| FRS-034 | Integrate preventive maintenance tasks with CMMS and preserve validated configuration. | Ensure maintenance performed without unauthorized changes to validated state. | PM tasks scheduled/executed in CMMS; PMs do not alter validated parameters without change control. | Execute PM task and verify record creation and no unauthorized parameter changes. | Y | Maintenance; QA |
| FRS-035 | Track calibration intervals and store traceable calibration certificates for instruments and sensors. | Ensure measurement accuracy and traceability to standards. | Calibration records attached to instrument IDs with due dates and alerts for upcoming calibrations. | Review instrument records and verify calibration certs and due date alerts. | Y | Validation Lead; QA |
| FRS-036 | Provide standardized mounting procedures and photographic evidence capture for data logger placement. | Ensure repeatable mapping results and comparability across runs. | Mounting SOPs available and photographic evidence required for each mapping run. | Perform mapping and verify photos and mounting SOP adherence in record. | Y | Validation Lead; Engineering |
| FRS-037 | Detect utility failures (steam loss/low pressure) and notify operators per SOP with logging. | Prevent compromised cycles and enable rapid response. | Utility sensors trigger alarms, notifications routed to operators, and events logged. | Simulate utility failure and verify alarm, notification, and log entry. | Y | Facilities; Engineering; QA |
| FRS-038 | Meet IT security policies for network segmentation, authentication, and vulnerability management. | Protect data integrity and system availability. | System passes vulnerability scan and conforms to IT security checklist. | Run vulnerability scan and review remediation evidence; verify segmentation and authentication controls. | N | IT; Security |
| FRS-039 | Perform automated backups and verified restores of cycle data and system configuration. | Ensure recoverability of validated configuration and records. | Backups executed per schedule and restore test recovers data within RTO/RPO. | Execute backup and restore test and verify data integrity and recovery time. | Y | IT; QA |
| FRS-040 | Provide archival and retrieval functions for electronic and hardcopy validation records with retention indexing. | Support audits and long-term retrieval of validation evidence. | Archive index searchable; retrieval test returns complete record set within defined time. | Generate archive index and perform retrieval test for sample validation package. | Y | QA; Document |
| FRS-041 | Gate operator access to production cycles until training and competency recorded. | Prevent untrained personnel from executing validated processes. | System blocks cycle start for users without required training/competency records. | Attempt cycle start with untrained user (expect block); allow trained user and verify log. | Y | Training; QA |
| FRS-042 | Link SOPs to HMI and require SOP confirmation at critical steps. | Ensure operators follow approved procedures during critical operations. | HMI displays SOP link and requires operator confirmation before critical steps; confirmations logged. | Execute critical step and verify HMI SOP prompt and confirmation logged. | Y | QA; Production |
| FRS-043 | Enforce BI/CI lot control, D-value, and expiry checks before use. | Prevent use of expired or non-conforming biological/chemical indicators. | System blocks use of expired/non-conforming lots and requires certificate attachment. | Attempt to use expired BI/CI (expect block); use valid lot and verify acceptance and attachment. | Y | Microbiology; QA |
| FRS-044 | Require positive and negative BI controls to be recorded and linked to each PQ run. | Validate incubation conditions and BI viability for each PQ. | PQ record includes positive/negative control entries and incubation results linked to run. | Record control BIs and verify incubation results and linkage in PQ record. | Y | Microbiology |
| FRS-045 | Capture full traceability for each load: device lot numbers, operator, cycle ID, BI/CI results. | Enable release decisions and investigations with complete traceability. | Load record schema contains required fields and supports barcode scanning; records retrievable. | Create load record with scanned device lots and verify fields populated and retrievable. | Y | Production; QA |
| FRS-046 | Ingest and store FAT/SAT deliverables and link them to DQ/IQ records. | Use supplier evidence to support qualification and reduce on-site scope. | FAT/SAT documents uploaded, indexed, and linked to equipment DQ/IQ records. | Upload FAT/SAT and verify linkage and accessibility from IQ record. | Y | Procurement; Validation Lead; QA |
| FRS-047 | Store material certificates and surface finish specs and link to equipment records. | Support cleaning/compatibility decisions and regulatory review. | Material certificates attached to equipment record and accessible during IQ/OQ/PQ. | Upload material certificates and verify retrieval and association with equipment. | Y | Procurement; Engineering; QA |
| FRS-048 | Record cleaning/decontamination events and require pre-PQ cleaning verification. | Prevent bioburden carryover and ensure clean starting condition for PQ. | Cleaning logs required and attached to PQ runs; pre-PQ cleaning verification completed. | Execute cleaning SOP, record entry, and verify attachment to PQ run. | Y | Production; QA |
| FRS-049 | Monitor ambient temperature and humidity and log values during cycles. | Ensure cycles run within acceptable ambient conditions that could affect performance. | Ambient sensors log T/H and system flags excursions that could affect cycle performance. | Record ambient T/H during cycles and verify logs and any flagged excursions. | N | Facilities; Engineering |
| FRS-050 | Enforce installation constraints for floor loading, anchoring, and clearances in IQ workflow. | Prevent improper installations that could affect performance or safety. | IQ checklist includes floor loading and anchoring verification and sign-off fields. | Complete IQ checklist and verify sign-offs and attached calculations. | N | Facilities; Engineering |
| FRS-051 | Verify condensate routing and trap function and log condensate events. | Prevent condensate backflow and support drying performance. | Condensate routing verified in IQ and condensate events logged during cycles. | Run cycle and observe condensate routing; verify trap function and log entries. | Y | Facilities |
| FRS-052 | Present audible/visual alarms and escalation workflow on HMI/operator consoles. | Ensure operator awareness and timely escalation of critical events. | HMI displays alarms, horn/strobe activates, and escalation notifications issued per SOP. | Trigger critical alarm and verify audible/visual alert and escalation notifications. | Y | Engineering; Production; QA |
| FRS-053 | Provide partial requalification templates and guided test scripts for minor repairs. | Reduce downtime while ensuring validated state after limited changes. | Partial requalification templates available and enforce required tests and sign-offs. | Execute partial requalification script after simulated minor repair and verify pass. | Y | Validation Lead; QA |
| FRS-054 | Export raw data files with provenance metadata and checksum to secure archive. | Preserve authenticity and integrity of validation evidence for audits. | Raw files exported with metadata and checksum and archived in secure repository. | Export raw data, verify metadata and checksum, and confirm archive entry. | Y | IT; Validation Lead |
| FRS-055 | Include measurement uncertainty fields in calibration records and apply uncertainty in acceptance logic. | Ensure correct interpretation of measurement results and acceptance decisions. | Calibration records include uncertainty and acceptance calculations incorporate uncertainty. | Review calibration certs and verify uncertainty used in acceptance calculations for mapping/pressure tests. | Y | Validation Lead; QA |
| FRS-056 | Provide lumen test fixtures and BI insertion workflows for lumen penetration testing. | Ensure worst-case lumen penetration is challenged and recorded. | Lumen test workflows available and BI insertion recorded with lumen dimensions. | Execute lumen penetration workflow and verify BI placement records and results. | Y | Microbiology; Validation Lead |
| FRS-057 | Provide Bowie-Dick or equivalent test program scaled to chamber size with pass/fail logic. | Verify air removal performance appropriate to chamber size and configuration. | BD program available, runs per standard, and records pass/fail in logs. | Run BD program and verify pass/fail output and log entry. | Y | Validation Lead; Engineering |
| FRS-058 | Integrate change control with requalification triggers and automated notifications. | Ensure requalification is initiated when changes could affect validated state. | Change control entries evaluate requalification impact and trigger workflows when criteria met. | Create change control entry and verify requalification trigger and notification routing. | Y | QA; Validation Lead |
| FRS-059 | Provide deviation and CAPA workflow with evidence attachment and closure verification. | Ensure systematic resolution and prevention of recurrence for deviations. | Deviation/CAPA records capture root cause, actions, verification, and closure evidence. | Create deviation, route through CAPA, and verify closure with evidence. | Y | QA |
| FRS-060 | Generate final validation report with traceability matrix, raw data attachments, BI/CI results, deviations, CAPA, and approvals. | Consolidate evidence for release and regulatory inspection. | Report generator produces complete report with attachments and signature fields. | Generate draft final report and verify inclusion of required sections and attachments. | Y | Validation Lead; |
| FRS-061 | Provide material compatibility test workflows and record pre/post cycle results for polymers, elastomers, adhesives, and coatings. | Ensure materials used in devices tolerate validated steam cycles without unacceptable degradation. | Test workflows exist for each material class; pre/post results recorded and pass/fail per defined limits. | Execute material compatibility tests (visual, dimensional, mechanical, chemical) and attach results to material record. | Y | R&D; Materials Engineering; QA |
| FRS-062 | Support packaging material matrix tests and record seal integrity, CI/BI outcomes, and peel test data. | Packaging behaves differently under steam/drying and must be validated per material/seal type. | Packaging matrix entries present; seal peel and CI/BI results recorded and meet acceptance criteria. | Run packaging PQ simulations, perform seal peel tests, and verify CI/BI and peel data recorded. | Y | Packaging Engineering; QA; Production |
| FRS-063 | Implement lumen/complex geometry matrix workflows capturing lumen length/ID and nested configurations with BI linkage. | Lumens and complex geometries are worst-case for steam penetration and require systematic challenge. | Lumen matrix populated; BI results linked to specific lumen dimensions and pass/fail recorded. | Execute lumen matrix tests with BIs at lumen tips and verify results recorded per matrix entry. | Y | Microbiology; Validation Lead |
| FRS-064 | Provide workflows to validate assembled and disassembled device states, including functional checks post-cycle. | Assembly state can change steam penetration and device function; both states must be validated if processed. | Separate workflows for assembled/disassembled states with BI/CI and functional test capture; pass criteria met. | Run both workflows and verify BI/CI and functional test results recorded and acceptable. | Y | Production; Validation Lead |
| FRS-065 | Support stability study planning, sample tracking, and periodic sterility/pack integrity checks for shelf-life claims. | Post-sterilization stability must be demonstrated for claimed shelf life and storage conditions. | Stability plan created; samples tracked; periodic sterility and packaging integrity results recorded. | Create stability plan, store samples under defined conditions, and record periodic test results. | Y | Regulatory Affairs; QA; R&D |
| FRS-066 | Provide water/condensate sampling workflows, result capture, and alarm logic for out-of-spec parameters. | Water quality affects steam purity, corrosion, and cycle performance. | Sampling events recorded; lab results attached; alarms generated for excursions per URS limits. | Perform water/condensate sampling, attach lab reports, and verify alarm behavior on excursions. | Y | Facilities; QA |
| FRS-067 | Apply altitude/ambient correction factors to cycle parameter calculations and document adjustments. | Atmospheric pressure at site elevation affects vacuum and lethality calculations. | Altitude correction fields present; adjusted parameters documented in cycle reports and DQ. | Configure altitude correction and verify adjusted cycle parameters and documentation during OQ. | Y | Validation Lead; Engineering |
| FRS-068 | Integrate HVAC/environmental monitoring for sterilized product handling and storage with data linkage to storage lots. | Post-sterilization environment can affect sterility and packaging integrity. | Environmental data linked to storage lot records; alerts generated for excursions. | Verify environmental data capture and linkage to storage lot records; simulate excursion and verify alert. | Y | Facilities; QA |
| FRS-069 | Record condensate/exhaust monitoring results and link to environmental compliance records and permits. | Environmental compliance for effluent and emissions must be demonstrable. | Condensate/exhaust sample results attached to environmental records and compared to permit limits. | Attach sample lab reports and verify compliance checks against permit thresholds. | N | Facilities; EHS; QA |
| FRS-070 | Enforce firmware/software patch management workflow requiring predeployment testing, rollback plan, and change control approval. | Patches can alter validated behavior; controlled deployment is required. | Patch cannot be applied to production without test evidence, rollback plan, and approved change control. | Attempt patch deployment without approval (expect block); deploy with approval in test environment and verify rollback. | N | IT; Validation Lead; QA |
| FRS-071 | Synchronize system clocks to approved NTP server and log periodic drift checks. | Accurate timestamps are essential for audit trails and traceability. | NTP configured; periodic drift checks logged and within defined tolerance. | Verify NTP configuration and review drift logs over defined interval. | N | IT; QA |
| FRS-072 | Encrypt data in transit and at rest per IT policy and log secure transfer events. | Protect data integrity and confidentiality for cycle and validation records. | TLS and disk encryption enabled; secure transfer logs present and verifiable. | Verify encryption settings and perform secure transfer test with log verification. | N | IT; Security |
| FRS-073 | Provide API endpoints, data mapping, and reconciliation checks for MES/LIMS integration. | Ensure accurate enterprise data exchange for lot release and traceability. | API endpoints available; mapping documented; reconciliation tests pass. | Execute integration test scripts and verify data reconciliation between systems. | N | IT; Production; QA |
| FRS-074 | Implement backup architecture meeting defined RTO/RPO and support restore verification tests. | Ensure recoverability of validated configuration and records after outages. | Backup/restore tests meet RTO/RPO and restore integrity verified. | Perform restore test and measure recovery time and data integrity. | N | IT; QA |
| FRS-075 | Enforce lot release rules and automated hold/release logic tied to BI/CI results and approvals. | Prevent premature release of product and ensure microbiology/QA sign-off. | System holds loads until BI/CI pass and approvals recorded; automated release executes on pass. | Simulate BI fail (expect hold) and BI pass (expect release) and verify workflow. | Y | QA; Production; IT |
| FRS-076 | Provide statistical sampling plan templates and acceptance calculation tools for batch releases. | Statistically justified sampling supports defensible release decisions. | Sampling templates available; acceptance calculations performed and recorded per plan. | Execute sampling plan on representative batch and verify recorded results and acceptance decision. | Y | QA; Biostatistics |
| FRS-077 | Enforce indicator selection rules (BI species, CI classes) and block use of non-conforming lots. | Ensure appropriate challenge and valid immediate verification. | Indicator master data present; system blocks expired/non-conforming lots and requires certificate attachment. | Attempt to select non-conforming indicator (expect block); select conforming indicator and verify acceptance. | Y | Microbiology; QA |
| FRS-078 | Capture environmental monitoring data for sterilized product storage areas and generate trend reports. | Ongoing monitoring preserves sterility and packaging integrity during storage. | Environmental data logged and trend reports generated per cadence with alerting for excursions. | Simulate environmental excursion and verify alert, log, and trend report generation. | N | QA; Facilities |
| FRS-079 | Provide label printing and electronic traceability fields for load and pack records (barcode support). | Labels and electronic traceability enable product tracking and recall capability. | Labels print required fields; load/pack records include traceability fields and barcode linkage. | Print sample labels and verify data accuracy and linkage to load/pack records. | Y | Production; QA |
| FRS-080 | Track spare part lead times, minimum on-hand quantities, and generate reorder alerts. | Minimize downtime and ensure availability of critical components to preserve validated state. | Spare part records include lead times and min stock; reorder alerts generated when below threshold. | Deplete part to below min and verify reorder alert generation and notification. | N | Procurement; |
| FRS-081 | Flag maintenance actions that require partial or full requalification and route to change control. | Ensure maintenance that affects validated state triggers appropriate requalification. | Maintenance actions mapped to requalification triggers and change control workflow initiated when required. | Execute maintenance action that should trigger requalification and verify workflow initiation and notifications. | Y | Maintenance; QA; Validation Lead |
| FRS-082 | Manage consumable qualification records and block use of unqualified consumables. | Consumables affect performance and must be controlled to preserve validation. | Consumable master data includes qualification status; system blocks unqualified items from use. | Attempt to use unqualified consumable (expect block); use qualified consumable and verify acceptance. | Y | Procurement; QA |
| FRS-083 | Support decommissioning workflow and archive records per retention policy. | Ensure proper disposition and record keeping for retired validated equipment. | Decommissioning checklist executed and records archived with retention metadata. | Execute mock decommissioning and verify archive entries and retention metadata. | N | QA; Facilities |
| FRS-084 | Classify reportable events and provide regulatory notification workflow with timelines. | Ensure timely regulatory reporting for defined events (e.g., widespread BI failures). | Event classification rules present and notification workflow enforces timelines and templates. | Simulate reportable event and verify notification workflow and timeline adherence. | Y | Regulatory Affairs; QA |
| FRS-085 | Provide emergency SOP repository, drill scheduling, and drill logging. | Ensure readiness and documented response for critical failures. | Emergency SOPs accessible, drills scheduled, and drill logs captured with outcomes and corrective actions. | Conduct drill and verify log entries, notifications, and corrective action tracking. | Y | QA; Production; EHS |
| FRS-086 | Generate audit packs per template on demand and verify completeness. | Speed audit response and ensure consistent evidence presentation. | Audit pack generator produces complete dataset per template and retrieval time within SLA. | Generate audit pack and verify inclusion of required documents and retrieval time. | N | QA; Document Control |
| FRS-087 | Store local/regional regulatory permits and alert on expirations. | Maintain compliance with local environmental, waste, and safety rules. | Permit documents stored and expiration alerts generated per schedule. | Upload permit and simulate expiration alert; verify workflow and notifications. | N | Regulatory Affairs; EHS |
| FRS-088 | Manage change control for consumable/spare substitutions including impact assessment and requalification requirements. | Prevent unassessed substitutions that could affect validated state. | Change control captures alternative part, impact assessment, and required verification tests. | Submit substitution change and verify impact assessment and test requirements recorded. | Y | Procurement; QA; Engineering |
| FRS-089 | Schedule recurring training and competency requalification and gate access accordingly. | Maintain operator and maintenance competence over time and after changes. | Training schedule enforced, competency tests recorded, and access gated until completion. | Schedule training, run competency test, and verify access gating and records. | Y | Training; QA |
| FRS-090 | Produce trend reports for BI/CI failures, deviations, and maintenance events and trigger management review. | Detect systemic issues early and drive CAPA. | Trend reports generated per cadence and management review actions recorded. | Generate trend report and verify management review minutes and action items. | Y | QA; Validation Lead |
| FRS-091 | Support limited validation workflows and risk assessment for emergency/alternative cycles. | Allow controlled use of emergency cycles with documented risk and limited validation. | Emergency cycle workflow includes risk assessment, limited OQ/PQ tests, and approval fields. | Execute emergency cycle workflow in test environment and verify required approvals and test evidence. | N | Validation Lead; QA |
| FRS-092 | Record third-party service provider qualifications and require approved procedures for work on critical components. | Ensure external work preserves validated state and traceability. | Service provider records present and work orders require provider qualification and sign-off. | Review sample service report and verify provider qualification and sign-off fields. | N | Maintenance; Procurement; QA |
| FRS-093 | Include accessories in load templates and validation workflows with compatibility checks. | Accessories affect heat transfer and must be validated as part of loads. | Accessories selectable in load templates and compatibility test results recorded. | Add accessory to load template and verify compatibility test results and mapping. | Y | Production; Validation Lead |
| FRS-094 | Provide nonconforming product workflow with quarantine, investigation, and disposition records. | Ensure consistent handling and investigation of suspect product. | Nonconforming events create quarantine flags and route disposition workflow with evidence capture. | Create nonconforming event and verify quarantine and disposition workflow execution. | Y | QA; Production |
| FRS-095 | Schedule and record steam trap and separator maintenance and verify steam quality post-maintenance. | Maintain steam quality critical to sterilization performance. | Maintenance tasks scheduled and post-maintenance steam sample results attached and acceptable. | Execute maintenance task and verify post-maintenance steam sample results attached. | Y | Facilities; Maintenance |
| FRS-096 | Provide MSA/Gage R&R execution and reporting tools for critical measurement instruments. | Ensure measurement systems are capable and uncertainties quantified. | MSA studies executed and results meet capability criteria; reports stored. | Conduct Gage R&R and verify report and acceptance. | Y | Validation Lead; QA |
| FRS-097 | Schedule annual VMP review and capture revision history and approvals. | Keep validation strategy current with process and regulatory changes. | Annual review scheduled; minutes recorded; VMP revisions tracked with approvals. | Execute VMP review and verify revision record and approvals. | N | QA; Validation Lead |
| FRS-098 | Store supplier quality agreements and track supplier deliverables (FAT/SAT, spares, software support). | Define supplier responsibilities that affect validated state. | Supplier QA agreements uploaded and deliverable checklist verified. | Review supplier agreement and confirm deliverables present and accepted. | N | Procurement; QA |
| FRS-099 | Provide environmental test workflows for seasonal/extreme ambient condition verification. | Ensure performance under ambient extremes or document mitigations. | Environmental test plan executed and results recorded or mitigations documented. | Run OQ tests under simulated extremes and verify results or mitigation evidence. | N | Facilities; Validation Lead |
| FRS-100 | Include measurement uncertainty propagation in acceptance calculation tools and reports. | Ensure acceptance decisions account for instrument uncertainty. | Acceptance calculations include uncertainty and are documented in OQ/PQ reports. | Perform acceptance calculation with uncertainty and verify documentation in reports. | Y | Validation Lead; QA |
| FRS-101 | Enforce pre-cleaning SOPs and record incoming bioburden sampling results. | Reduce pre-sterilization bioburden that could compromise lethality or BI interpretation. | Pre-cleaning records and bioburden sample results attached and within defined limits. | Execute pre-cleaning and bioburden sampling and verify records and results. | Y | Production; Microbiology; QA |
| FRS-102 | Maintain approved detergents list and record usage and compatibility test results. | Prevent residual cleaning agents from affecting sterilization or packaging. | Only approved detergents selectable; usage logged; compatibility evidence attached. | Attempt to select unapproved detergent (expect block); select approved and verify usage log. | N | R&D; QA; Production |
| FRS-103 | Support validation workflows for each steam source used (central plant/local generator) and link source ID to cycles. | Different steam sources can vary in quality and must be validated separately. | Steam source selectable and cycle records include source ID and sample results. | Run cycles using each source and verify source ID and steam sample linkage in records. | Y | Facilities; Validation Lead |
| FRS-104 | Ensure UPS and emergency power workflows support safe shutdown and data preservation. | Power interruptions can corrupt cycle data and create unsafe conditions. | UPS configured; failover tested; data preservation verified during simulated outage. | Simulate power loss and verify UPS behavior and data preservation. | N | IT; Facilities; Engineering |
| FRS-105 | Record EMC/grounding checks and attach EMC test reports to IQ records. | EMI can affect controller operation and data integrity. | EMC test results attached to IQ and grounding checks passed. | Review EMC documentation and perform site grounding check; attach results. | N | Engineering; IT |
| FRS-106 | Monitor noise and vibration during operation and log results against limits. | Excessive vibration/noise can affect equipment and operator safety. | Noise/vibration measurements recorded and within occupational/equipment limits. | Measure noise/vibration during operation and verify logs and compliance. | N | Facilities; EHS |
| FRS-107 | Provide seasonal ambient test workflows or document mitigations where extremes apply. | Ambient extremes can affect drying and control loops; mitigations may be required. | Seasonal test results or mitigation evidence attached to OQ. | Execute seasonal test or document engineering mitigations and verify evidence. | N | Facilities; Validation Lead |
| FRS-108 | Provide MSA/Gage R&R execution and reporting tools for critical instruments (repeatability/reproducibility). | Ensure measurement systems meet capability criteria for validation data. | MSA studies executed and meet predefined capability criteria; reports stored. | Conduct MSA/Gage R&R and verify results and report storage. | Y | Validation Lead; QA |
| FRS-109 | Monitor sensor drift and trigger recalibration workflows when thresholds exceeded. | Sensor drift can invalidate mapping and cycle control; proactive detection required. | Drift thresholds configured and recalibration workflow triggered automatically. | Simulate drift and verify detection and recalibration trigger and records. | Y | Engineering; QA |
| FRS-110 | Provide firmware rollback and recovery procedures and test capability in non-production environment. | Ability to revert to validated firmware if an update causes issues. | Rollback procedure documented and tested without data loss in test environment. | Perform rollback test in non-production and verify data integrity and recovery. | N | IT; Supplier; QA |
| FRS-111 | Restrict and log remote vendor access sessions with approval workflow and time limits. | Remote access can introduce security and change risks; control required. | Remote sessions require approval, are time-limited, and session logs stored. | Simulate vendor remote session and verify approval, time limit, and logs. | N | IT; QA; Supplier |
| FRS-112 | Enforce mobile device access policies and capture mobile action audit logs. | Mobile access can affect data security and operator behavior; must be controlled. | Mobile access requires authentication and actions logged in audit trail. | Perform mobile login and action and verify audit trail entry. | N | IT; QA |
| FRS-113 | Integrate release interlocks with MES/LIMS and prevent automated release until BI/CI and approvals complete. | Automated release must not bypass microbiology or QA checks. | MES/LIMS receives release status only after BI/CI pass and approvals; reconciliation passes. | Simulate BI fail and verify no release to MES/LIMS; simulate pass and verify release and reconciliation. | Y | IT; QA; Production |
| FRS-114 | Track retained samples with storage location, retention duration, and retrieval logs. | Retained samples support investigations and stability testing. | Retained sample records include location, retention period, and retrieval history. | Create retention record and perform retrieval; verify logs and sample integrity. | N | QA; Production |
| FRS-115 | Require post-maintenance verification tests and sign-off before return to production. | Major maintenance can alter validated performance and must be verified. | Post-maintenance verification script executed and pass required before production flag cleared. | Perform maintenance, execute verification script, and verify pass and sign-off. | Y | Maintenance; Validation Lead; QA |
| FRS-116 | Maintain obsolescence register and alternative part approval workflow with impact assessment. | Supplier part obsolescence can affect long-term maintainability and validated state. | Obsolescence entries present and alternative parts approved with impact assessments. | Review obsolescence register and verify alternative approval records and impact assessments. | N | Procurement; Engineering; QA |
| FRS-117 | Validate orientation and stacking templates and enforce their use in production. | Orientation and stacking affect steam penetration and drying; must be controlled. | Orientation/stacking templates enforced and recorded in load records; BI/CI acceptance met. | Execute orientation/stacking test and verify template enforcement and BI/CI results. | Y | Production; Validation Lead |
| FRS-118 | Provide mixed-load templates and BI placement strategies for mixed device processing. | Mixed loads create complex worst-case scenarios and require defined strategies. | Mixed-load templates available and BI/CI placement recorded per strategy; acceptance met. | Run mixed-load test and verify BI/CI results and template usage. | Y | Production; Validation Lead |
| FRS-119 | Apply enhanced handling, documentation, and release controls for implantable device loads. | Implants may have stricter release and documentation requirements. | Implant workflows require enhanced documentation and regulatory checklist sign-off. | Execute implant workflow and verify enhanced documentation and approvals. | Y | Regulatory Affairs; QA; Production |
| FRS-120 | Provide endotoxin control workflows and testing capture where applicable. | Steam sterilization does not remove endotoxin; controls must prevent introduction. | Endotoxin test procedures available and results linked to lot records where required. | Execute endotoxin test workflow and verify results attached to lot record. | Y | QA; R&D; Microbiology |
| FRS-121 | Require post-move IQ/OQ workflows and transport inspection capture after relocation. | Relocation can affect alignment, seals, and sensors; post-move verification required. | Post-move IQ/OQ scripts executed and pass before production; transport inspection recorded. | Perform post-move IQ/OQ and verify inspection records and pass criteria. | Y | Validation Lead; Engineering |
| FRS-122 | Provide cross-site comparability data capture and harmonized protocol templates for multi-site deployments. | Ensure consistent validation approach and comparability across sites using same model. | Cross-site DQ/OQ comparability fields present and harmonized protocol templates available. | Compare site OQ data to reference site and document comparability assessment. | N | Corporate Validation; QA |
| FRS-123 | Track supplier audit schedules and store audit reports and corrective actions. | Supplier quality affects validated state and long-term support. | Supplier audit records and CAPA follow-up stored and searchable. | Review supplier audit record and verify CAPA closure evidence. | N | Procurement; QA |
| FRS-124 | Include accessory cleaning and compatibility tests in validation workflows and PQ load patterns. | Accessories may retain moisture or bioburden and affect cycles. | Accessory cleaning validation records and PQ inclusion evidence present. | Execute accessory cleaning validation and verify PQ inclusion and results. | Y | Production; Validation Lead |
| FRS-125 | Capture provenance metadata and chain of custody for raw data files. | Ensure authenticity and integrity of validation evidence for audits. | Raw files include metadata (creator, timestamp, checksum) and chain of custody entries. | Export raw file and verify metadata and chain of custody entries in archive. | Y | IT; QA |
| FRS-126 | Provide redaction workflows that preserve validation evidence when sharing records externally. | Protect confidential information while retaining necessary validation evidence. | Redaction tool produces redacted records with audit trail and retains validation evidence. | Redact sample record and verify retained evidence and approval log. | N | QA; Regulatory Affairs; Legal |
| FRS-127 | Support water hardness testing workflows and link mitigation controls to cycle acceptance. | Water hardness affects steam generation and scale formation; mitigations may be required. | Water hardness results recorded and mitigation controls applied when outside limits; cycle acceptance documented. | Perform water hardness test and verify mitigation controls and record linkage. | N | Facilities; Validation Lead |
| FRS-128 | Schedule and record trap/separator maintenance and verify steam quality after maintenance. | Proper trap/separator function is critical to steam quality and cycle performance. | Maintenance records and post-maintenance steam sample results attached and acceptable. | Execute maintenance and verify post-maintenance steam sample results. | Y | Facilities; Maintenance |
| FRS-129 | Provide humidity control/monitoring and drying adjustment capabilities for drying phase. | Ambient and chamber humidity affect drying efficiency and residual moisture. | Humidity logged and drying adjustments applied; residual moisture acceptance met across defined humidity range. | Run drying tests under varied humidity and verify residual moisture results. | N | Validation Lead; Facilities |
| FRS-130 | Provide functional test workflows for embedded electronics pre/post sterilization per vendor guidance. | Electronics may be sensitive to steam and temperature cycles and require functional verification. | Pre/post functional test results recorded and meet vendor acceptance criteria. | Execute functional tests before and after sterilization and verify results. | N | R&D; QA; Supplier |
| FRS-131 | Support CI class selection and interpretation SOPs and record CI results accordingly. | Different CI classes provide different information and must be interpreted correctly. | CI class selectable per load type and interpretation recorded per SOP in PQ records. | Run CI class test and verify interpretation recorded and linked to PQ record. | Y | Microbiology; QA |
| FRS-132 | Provide alternative cycle recipe management and material compatibility test capture for special materials. | Some materials require non-standard cycles that must be validated and controlled. | Alternative recipes stored with material compatibility evidence and change control approval. | Create alternative recipe, attach material test evidence, and verify approval workflow. | Y | Validation Lead; R&D; QA |
| FRS-133 | Synchronize controller clocks and log periodic timestamp verification including DST/leap second handling. | Timestamp integrity is critical for audit trails and traceability. | Timestamps synchronized and periodic verification logs present and within tolerance. | Verify NTP sync and review timestamp verification logs. | N | IT; QA |
| FRS-134 | Manage door seal part numbers and replacement workflows including leak test verification. | Seal material affects vacuum integrity and leak rates; replacements must be verified. | Seal replacement workflow requires leak test and records part/lot numbers and verification results. | Replace seal per workflow and perform leak test; verify records and pass criteria. | Y | Maintenance; QA |
| FRS-135 | Provide carrier type selection and mapping templates for different thermal mass carriers. | Thermal mass affects heating/cooling profiles and may change lethality; carriers must be validated. | Carrier types selectable and mapping templates applied; mapping results recorded per carrier. | Run mapping with each carrier type and verify results and template usage. | Y | Production; Validation Lead |
| FRS-136 | Capture pre/post chemical analysis workflows for adhesives, labels, and inks when present. | Adhesives and inks can migrate or degrade, affecting product quality. | Pre/post analysis results recorded and compared to acceptance criteria; no unacceptable migration observed. | Perform chemical analysis and verify results and pass/fail documentation. | N | R&D; QA |
| FRS-137 | Provide nested packaging templates and BI placement workflows for nested cores. | Nested configurations can impede steam penetration and drying and must be challenged. | Nested templates available and BI results for nested cores recorded and pass/fail flagged. | Execute nested configuration test and verify BI/CI results and records. | Y | Production; Validation Lead |
| FRS-138 | Provide initial moisture measurement and pre-conditioning workflows and record results. | Initial moisture affects drying performance and may require pre-conditioning. | Initial moisture recorded and pre-conditioning applied when required; results logged and within acceptance. | Measure initial moisture, apply pre-conditioning, and verify post-cycle moisture results. | N | Validation Lead; Production |
| FRS-139 | Provide containment and biosafety workflows for loads containing biological materials or tissues. | Biological materials present unique contamination and safety considerations requiring containment. | Containment procedures enforced, biosafety approvals recorded, and specialized PQ scripts executed. | Execute specialized PQ with containment and verify biosafety controls and approvals. | Y | R&D; QA; EHS |
| FRS-140 | Capture lessons learned and support VMP/URS updates via management review workflows. | Continuous improvement ensures validation lifecycle evolves with field experience. | Management review records and VMP/URS revisions captured with action tracking and closure evidence. | Conduct management review and verify VMP/URS updates and action closure evidence. | N | QA; Validation Lead; |

Technical Requirement Specification TRS (Map to IQ Protocol)

| ID | TRS (Technical Requirement) | Rationale | Acceptance Criterion | Test Method | Critical to Quality | Responsible |
| --- | --- | --- | --- | --- | --- | --- |
| TRS-001 | Provide chamber, racks, fixtures sized and rated to accept wrapped instruments, trays, porous loads, rigid containers, and lumened devices. | Ensure physical accommodation of all validated load types. | Chamber and fixtures match DQ drawings and support specified load types; fit verification. | IQ dimensional inspection and photographic evidence; fit test with representative loads. | Y | Engineering; Supplier |
| TRS-002 | Supply adjustable racks, packaging fixtures, and drying hardware compatible with single/double wrap, rigid containers, and peel pouches. | Enable correct positioning and drying for each packaging format. | Fixtures installed and adjustable; drying hardware meets vendor spec. | IQ verify fixture installation; functional test cycles for each packaging format. | Y | Engineering; Supplier |
| TRS-003 | Chamber structural rating and cart interface to support gross load mass ≤50 kg; include load sensors if specified. | Ensure mechanical integrity and safe handling at max load. | Structural certs and load test results; cart interface tolerances met. | IQ structural certificate review; static load test with dummy mass. | Y | Engineering; Facilities |
| TRS-004 | Provide chamber internal dimensions, rack mounting points, and CAD drawings per DQ. | Ensure installation and rack compatibility per design. | As-built dimensions within ± specified tolerance of DQ drawings. | IQ dimensional verification and drawing sign-off. | N | Engineering |
| TRS-005 | Controller, heaters, valves, and vacuum pump sized and configured to execute pre-vacuum recipes (134°C/3 min; 132°C/4 min). | Provide required thermal and vacuum capability for lethality. | Controller executes recipes reproducibly; hardware meets power and capacity specs. | IQ verify hardware specs; OQ execute recipes and verify setpoints and plateaus. | Y | Engineering; Supplier |
| TRS-006 | Include Bowie-Dick program, vacuum profile control, and test log output capability. | Provide technical means to perform air removal diagnostics. | BD program present; vacuum profile configurable; test logs generated. | IQ verify software module; OQ run BD program and review logs. | Y | Engineering; Supplier |
| TRS-007 | Implement RBAC module, parameter limit tables, and secure audit logging for recipe edits. | Prevent unauthorized parameter changes at technical level. | RBAC enforces roles; parameter limits enforced; audit entries recorded. | IQ verify RBAC config; OQ attempt unauthorized edits and verify block and audit. | Y | IT; Supplier |
| TRS-008 | Controller hardware and sensor selection to achieve setpoint accuracy ±0.5°C; include calibration ports. | Ensure technical capability for accurate temperature control. | Sensor specs and controller accuracy meet ±0.5°C; calibration ports accessible. | IQ review sensor specs and certs; OQ compare controller vs reference probe. | Y | Engineering; QA |
| TRS-009 | Chamber steam distribution system, baffles, and sensor locations designed to achieve ±2.0°C uniformity. | Provide technical design for temperature uniformity. | CFD/engineering evidence and mapping results show ≤±2.0°C. | IQ review design; OQ temperature mapping in empty and loaded conditions. | Y | Engineering; Supplier |
| TRS-010 | Install redundant thermocouples at critical control points with discrepancy detection logic. | Provide failover and detection of sensor faults. | Redundant sensors installed; discrepancy alarm thresholds configured. | IQ verify sensor installation; OQ simulate discrepancy and verify alarm. | Y | Engineering |
| TRS-011 | Install calibrated pressure and vacuum transducers with accuracy ≤±2% FS and accessible calibration ports. | Ensure accurate pressure/vacuum measurement for control and leak detection. | Transducer certs traceable; readings within ±2% FS vs reference. | IQ review certs; OQ compare readings to calibrated gauges. | Y | Engineering; QA |
| TRS-012 | Vacuum pump, valves, seals, and chamber sealing designed to achieve ≤50 mbar and leak rate ≤5 mbar/min. | Provide technical vacuum capability and leak tightness. | Vacuum and leak rate achieved in OQ; hardware meets vendor spec. | IQ verify pump and seals; OQ vacuum hold and leak rate test. | Y | Engineering; Maintenance |
| TRS-013 | Steam separator, traps, and sampling ports installed; steam quality instrumentation provisioned. | Enable steam dryness and NCG monitoring and sampling. | Sampling ports accessible; separator/trap per vendor spec; instrumentation installed. | IQ inspect steam system; OQ sample steam and lab analysis for dryness/NCG. | Y | Facilities; Engineering |
| TRS-014 | Steam inlet piping, pressure regulator, gauges, and isolation valves sized for 3–6 bar operation. | Ensure stable steam supply and measurement. | Inlet pressure maintained 3–6 bar during cycles; hardware meets spec. | IQ piping inspection; OQ monitor inlet pressure during cycles. | Y | Facilities |
| TRS-015 | Controller storage, export capability, secure archive, and backup system configured for ≥10-year retention. | Provide technical data capture and long-term retention. | Export and archive tested; backups scheduled and verified. | IQ verify storage config; OQ export and restore test; archive retrieval test. | Y | IT; Supplier |
| TRS-016 | Audit database with immutable storage, secure timestamps, and search capability. | Ensure technical immutability and traceability of audit logs. | Audit entries immutable; timestamps synchronized; searchable. | IQ verify audit DB config; OQ create audit entries and attempt modification (expect fail). | Y | IT |
| TRS-017 | User management module with unique user IDs, password policy enforcement, and account lockout. | Provide technical enforcement of access controls. | Password policy active; lockout thresholds enforced; provisioning logs present. | IQ verify user module; OQ test password rules and lockout behavior. | Y | IT |
| TRS-018 | Electronic signature module with cryptographic linkage to records and signature metadata capture. | Enable secure e-signatures technically linked to records. | E-signature metadata captured and linked; cryptographic integrity verified. | IQ verify e-signature module; OQ apply e-signature and verify linkage. | N | IT; QA |
| TRS-019 | Export routines producing CSV/PDF with checksum/digital signature and secure transfer protocols. | Ensure exported data integrity and secure transfer. | Export files include checksum/signature; transfer logs recorded. | IQ verify export routines; OQ export sample and verify checksum/signature. | N | IT |
| TRS-020 | Alarm server, priority mapping, HMI alarm pages, horn/strobe outputs, and SOP link integration. | Provide technical alarm presentation and escalation mechanisms. | Alarms displayed with priority; horn/strobe operate; SOP links accessible. | IQ verify alarm hardware and HMI pages; OQ simulate alarms and verify outputs and SOP links. | Y | Engineering; Supplier |
| TRS-021 | Door interlock hardware, door position sensors, and interlock wiring with logging interface. | Provide technical interlock to prevent unsafe opening. | Interlock prevents opening under unsafe conditions; attempts logged. | IQ verify interlock wiring; OQ simulate unsafe condition and verify lock and log. | Y | Engineering; Safety |
| TRS-022 | E-stop circuit, UPS for controller, safe-state logic, and data flush on power loss. | Ensure safe shutdown and data preservation during power events. | E-stop and UPS function per spec; data preserved after simulated power loss. | IQ verify E-stop and UPS wiring; OQ simulate power loss and verify safe state and data integrity. | Y | Engineering; IT |
| TRS-023 | Firmware version control, firmware image repository, and V&V document linkage fields. | Ensure only validated firmware is installed and traceable. | Firmware version recorded and linked to V&V package; rollback images stored. | IQ verify firmware repo and linkage; OQ confirm installed version matches V&V. | Y | IT; Supplier |
| TRS-024 | Software V&V deliverables, test harness, and Part 11 control implementations (where applicable). | Provide technical evidence of software validation and controls. | V&V artifacts present; Part 11 controls implemented and tested. | IQ review V&V deliverables; OQ execute software test scripts and Part 11 checks. | N | IT; QA |
| TRS-025 | SPC module, data export for SPC, and statistical reporting tools integrated with controller logs. | Provide technical capability for reproducibility monitoring and SPC. | SPC charts generated; alerts for out-of-control conditions produced. | IQ verify SPC module; OQ run repeated cycles and generate SPC charts. | Y | IT; Validation Lead |
| TRS-026 | Load map templates, BI placement templates, and BI lot tracking fields in system database. | Enable technical recording of BI placement and traceability. | Templates present and BI placement recorded with lot IDs. | IQ verify templates; OQ execute BI placement workflow and verify records. | Y | IT; Microbiology |
| TRS-027 | CI placement templates, CI lot tracking, and CI result capture fields. | Provide technical capture of CI usage and results. | CI results recorded and flagged; lot certificates attachable. | IQ verify CI fields; OQ place CIs and verify result capture. | Y | IT; Production |
| TRS-028 | Drying control algorithm, exhaust control hardware, and moisture test ports for residual moisture measurement. | Provide technical means to control and verify drying performance. | Drying algorithm configurable; moisture ports accessible; residual moisture tests supported. | IQ verify hardware and ports; OQ run drying cycles and perform moisture tests. | Y | Engineering; Supplier |
| TRS-029 | Drying control repeatability features and logging for reproducibility analysis. | Ensure technical reproducibility of drying across cycles. | Drying logs show repeatable profiles and acceptable SD across runs. | OQ repeated drying cycles and statistical analysis of moisture results. | Y | Engineering; Validation Lead |
| TRS-030 | Load fixture drawings, mapping coordinates, and load template enforcement in HMI. | Ensure technical enforcement of load configurations and mapping consistency. | Load template selection required; mapping coordinates stored in cycle record. | IQ verify fixture drawings; OQ select template and verify stored coordinates. | Y | Engineering; IT |
| TRS-031 | PQ workflow module to sequence three consecutive production loads and enforce gating logic. | Provide technical enforcement of PQ sequencing and gating. | PQ workflow enforces sequence and prevents release until all passes. | IQ verify PQ module; PQ execute workflow and verify gating. | Y | IT; Validation Lead |
| TRS-032 | Change control integration, trigger rules, and requalification scheduling module. | Automate technical detection of changes requiring requalification. | Change control triggers requalification workflow and schedules tasks. | IQ verify integration; OQ create change and verify trigger and scheduling. | Y | IT; QA |
| TRS-033 | Spare parts database, part numbers, min stock fields, and reorder alert logic. | Provide technical spare tracking and alerting to minimize downtime. | Spare records present and reorder alerts generated when below min stock. | IQ verify database; OQ simulate depletion and verify alert. | Y | Procurement; Maintenance |
| TRS-034 | CMMS integration, PM task templates, and PM record capture without altering validated parameters. | Ensure PM execution is recorded and does not change validated config. | PM tasks recorded in CMMS; PMs cannot change validated parameters without change control. | IQ verify CMMS integration; OQ execute PM and verify records and parameter protection. | Y | Maintenance; IT |
| TRS-035 | Calibration certificate repository, instrument IDs, due date alerts, and calibration data fields including uncertainty. | Provide technical traceability and uncertainty capture for instruments. | Calibration certs attached to instrument IDs; alerts for upcoming calibrations. | IQ verify cert repository; OQ review certs and uncertainty fields. | Y | QA; Calibration Lab |
| TRS-036 | Mounting fixtures, mounting drawings, and photographic capture interface for data logger placement. | Ensure technical reproducibility of mapping sensor placement. | Mounting fixtures documented and photos required for each mapping run. | IQ verify fixtures; OQ perform mapping and verify photos attached. | Y | Engineering; Validation Lead |
| TRS-037 | Utility sensors (steam pressure, flow), alarm mapping, and notification routing configured. | Provide technical detection and notification for utility failures. | Utility sensors trigger alarms and notifications per SOP; events logged. | IQ verify sensor installation; OQ simulate utility failure and verify notifications. | Y | Facilities; Engineering |
| TRS-038 | Network segmentation, firewall rules, authentication mechanisms, and vulnerability scanning capability. | Provide technical cybersecurity controls for connected controllers. | Network meets IT security checklist and passes vulnerability scan. | IQ verify network config; OQ run vulnerability scan and review results. | N | IT; Security |
| TRS-039 | Backup architecture, backup schedule, restore procedures, and checksum verification for restored data. | Ensure technical recoverability of controller data and configuration. | Backups executed and restores verified within RTO/RPO; checksum matches. | IQ verify backup config; OQ perform restore and checksum verification. | Y | IT |
| TRS-040 | Archive repository with indexing, search, and retrieval performance meeting SLA. | Provide technical archival and retrieval for validation records. | Archive searchable and retrieval time within SLA; index verified. | IQ verify archive setup; OQ perform retrieval test for sample validation package. | Y | QA; IT |
| TRS-041 | Training record interface, gating logic to block cycle start for untrained users, and competency test capture. | Technically prevent untrained users from operating equipment. | System blocks cycle start for users without training records; competency results stored. | IQ verify training interface; OQ attempt cycle start with untrained user (expect block). | Y | Training; IT |
| TRS-042 | HMI links to SOP repository and confirmation prompt implementation for critical steps. | Provide technical enforcement of SOP adherence at operator interface. | HMI displays SOP link and requires confirmation before critical actions; confirmations logged. | IQ verify HMI links; OQ perform critical step and verify confirmation logged. | Y | IT; QA |
| TRS-043 | BI/CI lot fields, D-value fields, expiry checks, and certificate attachment capability. | Provide technical controls to prevent use of invalid indicators. | System blocks expired/non-conforming lots and requires certificate attachment. | IQ verify fields; OQ attempt to use expired lot (expect block) and use valid lot. | Y | IT; Microbiology |
| TRS-044 | Incubator integration fields for positive/negative controls and linkage to PQ records. | Ensure technical capture of BI control results linked to PQ runs. | Control BI results recorded and linked to PQ run record. | IQ verify incubator interface; OQ record control BI results and verify linkage. | Y | Microbiology; IT |
| TRS-045 | Load record schema with barcode/scan capability and database fields for device lot, operator, cycle ID, BI/CI results. | Provide technical traceability for each sterilized load. | Load records capture required fields and support barcode scanning. | IQ verify schema and scanner integration; OQ scan sample lots and verify record population. | Y | IT; Production |
| TRS-046 | Document repository for FAT/SAT deliverables with DQ/IQ linkage fields and version control. | Provide technical storage and linkage of supplier test evidence. | FAT/SAT documents uploaded and linked to DQ/IQ records with version control. | IQ verify repository and linkage; OQ review FAT/SAT attachments. | Y | Procurement; IT |
| TRS-047 | Material certificate upload fields, surface finish spec fields, and equipment record linkage. | Provide technical documentation for materials of construction. | Material certificates attached to equipment records and retrievable. | IQ upload certificates and verify retrieval from equipment record. | Y | Procurement; QA |
| TRS-048 | Cleaning log interface, cleaning SOP repository, and pre-PQ cleaning verification checklist. | Provide technical capture of cleaning events and verification prior to PQ. | Cleaning logs required and attached to PQ runs; pre-PQ verification completed. | IQ verify cleaning interface; OQ execute cleaning and verify log attachment. | Y | Production; QA |
| TRS-049 | Ambient T/H sensors, logging interface, and excursion flagging logic. | Provide technical monitoring of environmental conditions that may affect cycles. | Ambient sensors log T/H and system flags excursions per thresholds. | IQ verify sensor installation; OQ record ambient T/H during cycles and verify flags. | N | Facilities; IT |
| TRS-050 | Installation checklist module capturing floor loading, anchoring, and clearance sign-offs with attachments. | Provide technical enforcement of installation constraints during IQ. | IQ checklist completed with attachments and sign-offs. | IQ complete installation checklist and verify attachments (calculations, photos). | N | Facilities; Engineering |
| TRS-051 | Drain piping, trap configuration, condensate return, and inspection ports with monitoring fields. | Provide technical means to route and monitor condensate. | Drain routing verified and condensate events logged during cycles. | IQ inspect drain routing; OQ run cycle and observe condensate routing and logs. | Y | Facilities |
| TRS-052 | HMI alarm pages, horn/strobe outputs, escalation routing, and SOP link integration. | Provide technical alarm presentation and escalation mechanisms. | Alarms displayed and audible/visual outputs operate; escalation routing verified. | IQ verify alarm hardware and HMI; OQ simulate alarm and verify escalation. | Y | Engineering; IT |
| TRS-053 | Requalification templates, guided test scripts, and change control linkage for partial requalification. | Provide technical support for partial requalification after minor repairs. | Partial requalification templates present and enforce required tests and sign-offs. | IQ verify templates; OQ execute partial requalification script and verify pass. | Y | IT; Validation Lead |
| TRS-054 | Data export ports, metadata capture, checksum generation, and secure archive ingestion. | Ensure raw data provenance and integrity for archival. | Raw files exported with metadata and checksum and archived securely. | IQ verify export ports; OQ export raw files and verify metadata/checksum and archive entry. | Y | IT |
| TRS-055 | Calibration record fields including measurement uncertainty and automated inclusion in acceptance calculations. | Provide technical capture of uncertainty for acceptance decisions. | Calibration records include uncertainty and acceptance calculations incorporate uncertainty. | IQ verify calibration fields; OQ perform acceptance calc including uncertainty. | Y | QA; Calibration Lab |
| TRS-056 | Lumen test fixtures, BI insertion tools, and lumen dimension fields in test workflows. | Provide technical capability to challenge lumens and record dimensions. | Lumen fixtures available and BI insertion recorded with lumen dimensions. | IQ verify fixtures; OQ execute lumen penetration tests and verify records. | Y | Engineering; Microbiology |
| TRS-057 | BD pack compatibility list, BD test program, and pass/fail logic scaled to chamber size. | Provide technical BD testing appropriate to chamber configuration. | BD program runs and records pass/fail per defined criteria. | IQ verify BD program; OQ run BD tests and verify pass/fail logs. | Y | Validation Lead; Engineering |
| TRS-058 | Change control module with requalification trigger rules and automated notifications. | Provide technical automation of requalification triggers from change events. | Change control triggers requalification workflow and notifications sent. | IQ verify module; OQ create change and verify trigger and notifications. | Y | IT; QA |
| TRS-059 | Deviation/CAPA module with evidence attachment, root cause fields, and closure verification. | Provide technical CAPA management and evidence capture. | Deviation/CAPA records capture required fields and closure evidence stored. | IQ verify module; OQ create deviation and route through CAPA to closure. | Y | QA |
| TRS-060 | Report generator capable of compiling traceability matrix, raw data attachments, BI/CI results, deviations, CAPA, and signature fields. | Provide technical consolidation of validation evidence into final report. | Report generator produces complete report with attachments and signature fields. | IQ verify report generator; OQ generate draft report and verify contents. | Y | IT; Validation Lead |
| TRS-061 | Material test plan templates, sample handling fixtures, and result capture fields for pre/post cycle tests. | Provide technical support for material compatibility testing. | Test templates present and results captured with pass/fail flags. | IQ verify templates and fixtures; OQ execute material tests and verify results. | Y | R&D; Materials Engineering |
| TRS-062 | Packaging test fixtures, seal peel test equipment integration, and CI/BI placement templates. | Provide technical capability to validate packaging and seals. | Fixtures and equipment installed; test results recorded and meet acceptance. | IQ verify fixtures; OQ run packaging tests and record peel/CI/BI results. | Y | Packaging Engineering; QA |
| TRS-063 | Lumen matrix database, BI linkage fields, and nested configuration templates. | Provide technical management of lumen and complex geometry test matrix. | Lumen matrix populated and BI results linked to specific geometry entries. | IQ verify matrix; OQ execute matrix tests and verify linkage. | Y | IT; Microbiology |
| TRS-064 | Functional test fixtures and workflows for assembled and disassembled device states with result capture. | Provide technical means to validate both assembly states and device function. | Workflows capture BI/CI and functional test results for each state. | IQ verify fixtures; OQ run both workflows and verify results. | Y | Production; R&D |
| TRS-065 | Stability sample tracking module, storage condition fields, and scheduled test events. | Provide technical management of stability studies and sample traceability. | Stability plan created and samples tracked with scheduled tests and results recorded. | IQ verify tracking module; OQ create stability plan and verify scheduled events. | Y | Regulatory Affairs; QA |
| TRS-066 | Water sampling ports, conductivity/hardness sensors, lab result attachment fields, and alarm thresholds. | Provide technical monitoring of feedwater and condensate quality. | Sampling events recorded and alarms generated for out-of-spec results. | IQ verify ports and sensors; OQ perform sampling and verify lab result attachment and alarms. | Y | Facilities; QA |
| TRS-067 | Altitude correction parameter fields in controller and documentation fields for DQ adjustments. | Provide technical compensation for site elevation effects on cycles. | Altitude correction applied and documented in cycle reports; controller supports parameter adjustments. | IQ verify parameter fields; OQ run cycles with corrections and verify documentation. | Y | Engineering; Validation Lead |
| TRS-068 | HVAC interface points, environmental sensor integration, and data linkage to storage lot records. | Provide technical integration of environmental monitoring for storage areas. | Environmental data linked to storage lots and alerts generated for excursions. | IQ verify HVAC interfaces; OQ simulate excursion and verify linkage and alerts. | Y | Facilities; IT |
| TRS-069 | Condensate treatment/tank instrumentation, sampling ports, and permit linkage fields. | Provide technical capability to monitor and document environmental compliance. | Sample results attached and compared to permit thresholds; records retrievable. | IQ verify treatment/tank instrumentation; OQ attach sample reports and verify compliance checks. | N | Facilities; EHS |
| TRS-070 | Patch test environment, rollback image repository, and change control enforcement for production deployment. | Provide technical controls to manage software patches safely. | Patch cannot be deployed to production without approved change control and rollback plan. | IQ verify test environment and rollback repo; OQ attempt unauthorized patch (expect block) and authorized patch test. | N | IT |
| TRS-071 | NTP client configuration, time server settings, and timestamp drift logging. | Provide technical time synchronization for audit trail integrity. | NTP configured and drift logs within tolerance. | IQ verify NTP config; OQ monitor drift over defined interval. | N | IT |
| TRS-072 | TLS configuration, disk encryption, and secure transfer logging per IT policy. | Provide technical encryption for data in transit and at rest. | Encryption enabled and secure transfer logs present and verifiable. | IQ verify encryption settings; OQ perform secure transfer and verify logs. | N | IT; Security |
| TRS-073 | API endpoints, authentication tokens, data mapping spec, and reconciliation logs for MES/LIMS integration. | Provide technical integration points for enterprise systems. | API endpoints functional; mapping validated; reconciliation logs show matched records. | IQ verify API endpoints; OQ run integration tests and verify reconciliation. | N | IT |
| TRS-074 | Backup architecture (on-site/off-site), RTO/RPO settings, and automated restore scripts. | Provide technical disaster recovery capability for controller and data. | Restore tests meet RTO/RPO and data integrity verified. | IQ verify backup architecture; OQ perform restore and measure recovery time. | N | IT |
| TRS-075 | Release interlock logic, approval workflow engine, and MES/LIMS handoff API. | Provide technical enforcement of lot release rules and automated holds. | Interlock prevents release until BI/CI pass and approvals recorded; MES/LIMS receives release only after pass. | IQ verify interlock logic; OQ simulate BI fail/pass and verify hold/release and MES handoff. | Y | IT; QA |
| TRS-076 | Sampling plan templates, statistical calculation engine, and acceptance decision fields. | Provide technical support for statistical sampling and acceptance. | Sampling templates available and acceptance calculations performed and recorded. | IQ verify templates; OQ execute sampling plan and verify calculations and recorded decision. | Y | IT; QA |
| TRS-077 | Indicator master data (BI species, CI classes), lot control, and blocking logic for non-conforming lots. | Provide technical enforcement of indicator selection and lot validity. | Master data present and system blocks expired/non-conforming lots. | IQ verify master data; OQ attempt to use non-conforming lot (expect block). | Y | IT; Microbiology |
| TRS-078 | Environmental sensor network, logging, and trend report generator for storage areas. | Provide technical monitoring and trending of storage environment. | Environmental data logged and trend reports generated per cadence. | IQ verify sensors and logging; OQ generate trend report and verify contents. | N | Facilities; IT |
| TRS-079 | Label printer drivers, barcode format templates, and print verification interface. | Provide technical capability to print traceable labels and link to records. | Labels print required fields and barcodes scan to populate load records. | IQ verify printer integration; OQ print sample labels and scan to verify linkage. | Y | IT; Production |
| TRS-080 | Spare parts database with lead time fields, min stock, and automated reorder notification engine. | Provide technical spare management to reduce downtime. | Reorder alerts generated when stock below min and lead times recorded. | IQ verify database; OQ simulate depletion and verify alert generation. | N | Procurement; Maintenance |
| TRS-081 | Maintenance action catalog with mapped requalification triggers and automated change control initiation. | Provide technical mapping of maintenance to requalification requirements. | Maintenance actions mapped and change control initiated when triggers met. | IQ verify catalog; OQ execute maintenance action and verify trigger and change control initiation. | Y | Maintenance; IT |
| TRS-082 | Consumable master data, certificate attachment fields, and blocking logic for unqualified items. | Provide technical control to prevent use of unqualified consumables. | System blocks unqualified consumables and requires certificate attachment for qualified items. | IQ verify master data; OQ attempt to use unqualified consumable (expect block). | Y | Procurement; QA |
| TRS-083 | Decommissioning workflow module, archive export, and retention metadata capture. | Provide technical support for equipment retirement and record retention. | Decommissioning checklist executed and records archived with retention metadata. | IQ verify module; OQ execute mock decommissioning and verify archive entries. | N | QA; IT |
| TRS-084 | Event classification engine, notification templates, and regulatory timeline enforcement logic. | Provide technical capability to identify and manage reportable events. | Event classification triggers notification workflow and enforces timelines. | IQ verify engine; OQ simulate reportable event and verify notifications and timelines. | Y | Regulatory Affairs; IT |
| TRS-085 | Emergency SOP repository, drill scheduler, and drill log capture with corrective action linkage. | Provide technical readiness and documentation for emergency response. | SOPs accessible; drills scheduled and logs captured with corrective actions. | IQ verify repository; OQ run drill and verify log capture and corrective action entries. | Y | QA; IT |
| TRS-086 | Audit pack generator, index template, and retrieval performance metrics. | Provide technical generation of audit evidence packages on demand. | Audit pack generated per template and retrieval time within SLA. | IQ verify generator; OQ generate mock audit pack and verify completeness and retrieval time. | N | QA; IT |
| TRS-087 | Permit repository, expiration alert engine, and compliance checklist fields. | Provide technical tracking of local/regional regulatory permits. | Permits stored and expiration alerts generated per schedule. | IQ upload permit; OQ simulate expiration and verify alert workflow. | N | Regulatory Affairs; IT |
| TRS-088 | Change control module supporting alternative part entries, impact assessment fields, and verification test scheduling. | Provide technical management of consumable/spare substitutions and requalification impact. | Alternative parts recorded with impact assessment and verification tests scheduled. | IQ verify module; OQ submit substitution and verify impact assessment and test scheduling. | Y | IT; Procurement |
| TRS-089 | Training scheduler, competency test templates, and access gating integration. | Provide technical enforcement of recurring training and competency requalification. | Training scheduled and competency results recorded; access gated until completion. | IQ verify scheduler; OQ schedule training and verify gating behavior. | Y | Training; IT |
| TRS-090 | Trend analytics engine, report scheduler, and management review export fields. | Provide technical tools for trending BI/CI failures and deviations. | Trend reports generated per cadence and management review exports available. | IQ verify analytics engine; OQ generate trend report and verify export. | Y | QA; IT |
| TRS-091 | Emergency cycle template repository, risk assessment fields, and limited validation workflow. | Provide technical control for emergency/alternative cycle use and documentation. | Emergency templates present and limited validation workflow enforces required tests and approvals. | IQ verify templates; OQ execute emergency workflow in test environment and verify approvals. | N | Validation Lead; IT |
| TRS-092 | Supplier qualification records, service report templates, and signed work order capture fields. | Provide technical traceability for third-party service activities. | Supplier records present and service reports attached to equipment records. | IQ verify supplier records; OQ review sample service report and verify attachments. | N | Procurement; Maintenance |
| TRS-093 | Accessory master data, compatibility test fields, and inclusion in load templates. | Provide technical management of accessories used in loads. | Accessories selectable in templates and compatibility test results recorded. | IQ verify accessory data; OQ include accessory in load and verify compatibility test results. | Y | IT; Production |
| TRS-094 | Nonconforming product module with quarantine flags, investigation fields, and disposition records. | Provide technical workflow for handling suspect product consistently. | Nonconforming events create quarantine flags and route disposition workflow with evidence capture. | IQ verify module; OQ create nonconforming event and verify workflow execution. | Y | QA; IT |
| TRS-095 | Trap/separator maintenance schedule fields, sampling ports, and post-maintenance steam quality verification workflow. | Provide technical maintenance and verification for steam quality components. | Maintenance tasks scheduled and post-maintenance steam sample results attached and acceptable. | IQ verify schedule and ports; OQ perform maintenance and verify sample results. | Y | Facilities; Maintenance |
| TRS-096 | MSA/Gage R&R templates, data capture forms, and analysis tools integrated with validation reports. | Provide technical capability to perform and report measurement system analysis. | MSA studies executed and results meet capability criteria; reports stored and linked. | IQ verify templates; OQ conduct Gage R&R and verify report storage. | Y | QA; Validation Lead |
| TRS-097 | VMP review scheduler, version control, and management review minutes capture fields. | Provide technical support for periodic validation program review and record keeping. | Annual review scheduled and revision history captured with approvals. | IQ verify scheduler; OQ execute VMP review and verify revision capture. | N | QA |
| TRS-098 | Supplier QA agreement repository, deliverable checklist fields, and compliance tracking. | Provide technical storage and tracking of supplier obligations affecting validated state. | Supplier agreements uploaded and deliverable checklist verified. | IQ verify repository; OQ review supplier agreement and deliverable evidence. | N | Procurement; QA |
| TRS-099 | Environmental test harness and simulated ambient test setup capability for OQ. | Provide technical means to simulate seasonal/extreme ambient conditions during OQ. | Environmental test setup available and results recorded or mitigations documented. | IQ verify test harness; OQ run simulated ambient tests and record results. | N | Facilities; Validation Lead |
| TRS-100 | Uncertainty calculation templates and automated inclusion in acceptance decision logic and reports. | Provide technical propagation of measurement uncertainty into acceptance decisions. | Acceptance calculations include uncertainty and documented in OQ/PQ reports. | IQ verify templates; OQ perform acceptance calc including uncertainty and verify documentation. | Y | QA; Validation Lead |
| TRS-101 | Bioburden sampling procedure fields, lab result attachment, and acceptance threshold logic. | Provide technical capture and enforcement of pre-cleaning bioburden limits. | Bioburden sample results attached and compared to thresholds; system flags excursions. | IQ verify sampling fields; OQ perform bioburden sampling and verify result capture and flags. | Y | Microbiology; QA |
| TRS-102 | Approved detergents master list, compatibility test result fields, and usage logging. | Provide technical control of cleaning agents and evidence of compatibility. | Only approved detergents selectable and usage logged; compatibility evidence attached. | IQ verify master list; OQ attempt to select unapproved detergent (expect block) and verify usage log. | N | QA; Procurement |
| TRS-103 | Steam source identification fields, connection points, and sampling port mapping for each source. | Provide technical differentiation and validation capability for multiple steam sources. | Steam source ID recorded in cycle records and sample results linked to source. | IQ verify connection points; OQ run cycles using each source and verify source ID and sample linkage. | Y | Facilities; IT |
| TRS-104 | UPS sizing documentation, failover wiring, and power loss test scripts for controller and critical systems. | Provide technical resilience for safe shutdown and data preservation. | UPS supports safe shutdown and data preservation during simulated outage per spec. | IQ verify UPS wiring and sizing; OQ simulate power loss and verify data preservation. | N | IT; Facilities |
| TRS-105 | EMC test reports, grounding scheme drawings, and site grounding verification points. | Provide technical mitigation of EMI risks to controller operation. | EMC reports attached and site grounding verification passed. | IQ review EMC reports; OQ perform site grounding check and verify results. | N | Engineering; IT |
| TRS-106 | Noise and vibration measurement points, sensors, and logging interface for installation area. | Provide technical monitoring of noise/vibration to protect equipment and personnel. | Noise/vibration measurements recorded and within limits. | IQ verify sensor placement; OQ measure noise/vibration during operation and verify compliance. | N | Facilities; EHS |
| TRS-107 | Environmental test configuration files and mitigation documentation for seasonal extremes. | Provide technical evidence of performance or mitigations under ambient extremes. | Test results or mitigation evidence attached to OQ records. | IQ verify config; OQ execute seasonal tests or document mitigations and verify evidence. | N | Facilities; Validation Lead |
| TRS-108 | MSA/Gage R&R data capture templates, analysis scripts, and report export capability. | Provide technical tools to perform and report measurement system capability. | MSA studies executed and reports exported and stored. | IQ verify templates; OQ perform MSA and verify report export. | Y | QA; Validation Lead |
| TRS-109 | Drift monitoring algorithms, threshold fields, and automated recalibration workflow triggers. | Provide technical detection and response to sensor drift. | Drift detection triggers recalibration workflow automatically when thresholds exceeded. | IQ verify algorithm config; OQ simulate drift and verify trigger and recalibration record. | Y | Engineering; QA |
| TRS-110 | Firmware image repository, rollback scripts, and recovery verification procedures in test environment. | Provide technical rollback capability to validated firmware versions. | Rollback tested in non-production and data integrity preserved. | IQ verify repository; OQ perform rollback test in test environment and verify recovery. | N | IT; Supplier |
| TRS-111 | Remote access gateway, session approval workflow, session time limits, and session logging. | Provide technical control and auditability for vendor remote support. | Remote sessions require approval, time limits enforced, and session logs stored. | IQ verify gateway; OQ simulate remote session and verify approval and logs. | N | IT; Supplier |
| TRS-112 | Mobile access authentication, device management integration, and mobile action audit logging. | Provide technical control and audit trail for mobile operator interfaces. | Mobile access authenticated and actions logged in audit trail. | IQ verify device management integration; OQ perform mobile actions and verify audit logs. | N | IT |
| TRS-113 | API endpoints, interlock logic, and reconciliation reports for MES/LIMS release integration. | Provide technical enforcement of automated release interlocks and data handoff. | MES/LIMS receives release only after BI/CI pass and approvals; reconciliation reports match. | IQ verify API endpoints; OQ simulate BI fail/pass and verify MES handoff and reconciliation. | Y | IT; QA |
| TRS-114 | Sample retention fields, storage location mapping, and retrieval logging interface. | Provide technical tracking and retrieval of retained samples. | Retained sample records include location and retrieval history; retrieval test passes. | IQ verify retention fields; OQ create retention record and perform retrieval test. | N | QA; IT |
| TRS-115 | Post-maintenance verification scripts, hold flags, and sign-off fields to prevent production until pass. | Provide technical enforcement of post-maintenance verification before production. | Post-maintenance verification executed and pass required to clear production hold. | IQ verify scripts; OQ perform maintenance and execute verification script and sign-off. | Y | Maintenance; Validation Lead |
| TRS-116 | Obsolescence register, alternative part approval fields, and impact assessment templates. | Provide technical management of part lifecycle and approved alternatives. | Obsolescence entries present and alternative parts approved with impact assessments. | IQ verify register; OQ review alternative approval records and impact assessments. | N | Procurement; Engineering |
| TRS-117 | Orientation/stacking templates, fixture definitions, and BI placement maps for validated orientations. | Provide technical enforcement of validated orientations and stacking patterns. | Templates enforced and BI/CI acceptance met for validated orientations. | IQ verify templates; OQ run orientation tests and verify BI/CI results. | Y | Production; Validation Lead |
| TRS-118 | Mixed-load template definitions, BI placement strategy fields, and worst-case mapping support. | Provide technical support for validating mixed loads and worst-case coverage. | Mixed-load templates present and BI/CI placement recorded per strategy. | IQ verify templates; OQ run mixed-load tests and verify BI/CI results. | Y | Production; Validation Lead |
| TRS-119 | Implant handling workflow fields, enhanced traceability fields, and regulatory checklist integration. | Provide technical controls for implantable device handling and documentation. | Implant workflow enforces enhanced documentation and regulatory checklist sign-off. | IQ verify workflow; OQ execute implant workflow and verify documentation and approvals. | Y | Regulatory Affairs; IT |
| TRS-120 | Endotoxin test workflow fields, sample handling controls, and result linkage to lot records. | Provide technical capture of endotoxin testing where required. | Endotoxin results attached to lot records and acceptance logic applied. | IQ verify workflow; OQ perform endotoxin test and verify result linkage. | Y | QA; Microbiology |
| TRS-121 | Transport inspection checklist fields, post-move IQ/OQ scripts, and vibration measurement capture. | Provide technical verification after relocation to ensure equipment integrity. | Post-move IQ/OQ executed and vibration/inspection records attached and pass criteria met. | IQ verify checklist; OQ perform post-move IQ/OQ and verify vibration data. | Y | Validation Lead; Engineering |
| TRS-122 | Cross-site DQ/OQ comparability fields, FAT/SAT reference links, and harmonized protocol templates. | Provide technical comparability and harmonization across multiple sites. | Cross-site comparability data captured and harmonized protocols available. | IQ verify comparability fields; OQ compare site data to reference and document assessment. | N | Corporate Validation; IT |
| TRS-123 | Supplier audit record fields, audit schedule, and CAPA linkage for supplier corrective actions. | Provide technical tracking of supplier quality and follow-up. | Supplier audit records stored and CAPA follow-up evidence attached. | IQ verify audit record fields; OQ review supplier audit and CAPA evidence. | N | Procurement; QA |
| TRS-124 | Accessory cleaning SOP fields, compatibility test result capture, and PQ inclusion flags. | Provide technical inclusion of accessories in cleaning and validation workflows. | Accessory cleaning validation records present and PQ inclusion flagged. | IQ verify SOP fields; OQ execute accessory cleaning validation and verify PQ inclusion. | Y | Production; Validation Lead |
| TRS-125 | Raw data metadata capture (creator, timestamp, checksum), file hash generation, and chain of custody logs. | Provide technical provenance and integrity for raw validation data. | Raw files include metadata and checksum; chain of custody entries present. | IQ verify metadata capture; OQ export raw files and verify hash and chain of custody. | Y | IT; QA |
| TRS-126 | Redaction tool, redaction logs, and approval workflow preserving validation evidence. | Provide technical capability to redact sensitive info while retaining evidence. | Redacted records retain required validation evidence and redaction logs show approvals. | IQ verify tool; OQ redact sample record and verify retained evidence and approval log. | N | QA; IT |
| TRS-127 | Water hardness sampling ports, test kit integration fields, and mitigation control linkage to cycle acceptance. | Provide technical monitoring and mitigation for water hardness effects. | Hardness results recorded and mitigation controls applied when outside limits; cycle acceptance documented. | IQ verify ports; OQ perform hardness tests and verify mitigation linkage. | N | Facilities; QA |
| TRS-128 | Trap/separator maintenance schedule fields, sampling ports, and post-maintenance verification workflow. | Provide technical maintenance and verification for steam quality components. | Maintenance records and post-maintenance steam sample results attached and acceptable. | IQ verify schedule and ports; OQ perform maintenance and verify sample results. | Y | Facilities; Maintenance |
| TRS-129 | Humidity sensor integration, drying adjustment parameters, and logging for drying phase. | Provide technical control for humidity effects on drying performance. | Humidity logged and drying adjustments applied; residual moisture acceptance met across range. | IQ verify sensors; OQ run drying tests under varied humidity and verify results. | N | Facilities; Engineering |
| TRS-130 | Functional test fixtures, test scripts, and result capture for embedded electronics pre/post sterilization. | Provide technical verification of electronics functionality after cycles. | Pre/post functional test results recorded and meet vendor acceptance criteria. | IQ verify fixtures; OQ execute functional tests and verify results. | N | R&D; QA |
| TRS-131 | CI class master data, interpretation SOP fields, and CI result capture per class. | Provide technical support for correct CI interpretation per load type. | CI class selectable and interpretation recorded per SOP in PQ records. | IQ verify master data; OQ run CI class tests and verify interpretation capture. | Y | Microbiology; IT |
| TRS-132 | Alternative cycle recipe storage, recipe version control, and material compatibility evidence attachment. | Provide technical control for non-standard cycles and material compatibility. | Alternative recipes stored with attached material test evidence and change control approval. | IQ verify recipe storage; OQ create alternative recipe and verify attachments and approvals. | Y | IT; R&D |
| TRS-133 | NTP synchronization, timestamp verification logs, and DST/leap second handling logic. | Provide technical timestamp integrity across systems. | Timestamps synchronized and verification logs present; DST/leap second handled correctly. | IQ verify NTP config; OQ review timestamp logs and DST handling. | N | IT |
| TRS-134 | Door seal part number fields, replacement workflow, and automated leak test script execution. | Provide technical control for seal replacements and leak verification. | Replacement workflow requires leak test and records part/lot numbers and results. | IQ verify part fields; OQ replace seal and run leak test and verify records. | Y | Maintenance; QA |
| TRS-135 | Carrier type master data, mapping templates per carrier, and mapping point presets. | Provide technical mapping support for carriers with different thermal mass. | Carrier templates present and mapping results recorded per carrier type. | IQ verify carrier data; OQ run mapping for each carrier and verify results. | Y | Production; Validation Lead |
| TRS-136 | Chemical analysis integration fields, sample handling SOPs, and result capture for adhesives/inks. | Provide technical capability to detect migration or degradation of adhesives/inks. | Pre/post analysis results recorded and compared to acceptance criteria. | IQ verify analysis integration; OQ perform chemical analysis and verify results. | N | R&D; QA |
| TRS-137 | Nested load fixture definitions, BI placement templates for nested cores, and drying check fields. | Provide technical support to challenge nested configurations and verify drying. | Nested templates available and BI/CI results for nested cores recorded and pass/fail flagged. | IQ verify fixtures; OQ run nested tests and verify BI/CI and drying results. | Y | Production; Validation Lead |
| TRS-138 | Moisture measurement device integration, pre-conditioning workflow fields, and result logging. | Provide technical control for variable initial moisture content and pre-conditioning. | Initial moisture recorded and pre-conditioning applied when required; results logged. | IQ verify device integration; OQ measure initial moisture and verify pre-conditioning and post-cycle results. | N | Production; QA |
| TRS-139 | Containment fixture definitions, biosafety procedure fields, and specialized PQ script templates for biological materials. | Provide technical containment and biosafety controls for biological loads. | Containment enforced and specialized PQ scripts executed with biosafety approvals recorded. | IQ verify containment fixtures; OQ execute specialized PQ and verify biosafety controls. | Y | EHS; QA |
| TRS-140 | Management review module, lessons learned capture fields, and VMP/URS revision workflow. | Provide technical support for continuous improvement and VMP updates. | Management review records captured and VMP/URS revisions tracked with action closure evidence. | IQ verify module; OQ conduct management review and verify revision capture. | N | QA; |

Installation Qualification IQ Protocol

Purpose: Verify correct delivery, installation, utilities, documentation, safety devices, and readiness for OQ.

Prerequisites: DQ complete; FAT/SAT accepted and available; calibrated validation instruments on site; SOPs available; trained personnel scheduled; supplier V&V and firmware images available.

Execution Notes: Collect vendor FAT/SAT, wiring diagrams, calibration certificates, material certificates, photos of installation, and utility readings. Verify TRS items during each test. Document deviations immediately and route to CAPA if they affect sterility, safety, or validated configuration. Attach raw files, photos, and certificates to the IQ record.

IQ Protocol Table

| Test ID | Requirement (TRS mapping) | Test Method | Acceptance Criterion | Instrument Calibration | Frequency | Responsible |
| --- | --- | --- | --- | --- | --- | --- |
| IQ-01 | Equipment identity — TRS-001; TRS-023; TRS-005 | Verify model, serial, firmware, PO, and DQ package match physically and in system. | Model/serial/firmware/PO/DQ match; firmware image present in TRS repo. | PO packing list; firmware image printout; supplier V&V. | Once (at install) | Validation Lead |
| IQ-02 | Documentation delivered — TRS-046; TRS-024; TRS-047 | Confirm manuals, BOM, FAT/SAT, wiring diagrams, spare parts list, material certs, V&V deliverables present. | All documents present, approved, and linked to equipment record. | Document review checklist. | Once | QA |
| IQ-03 | Physical installation — TRS-004; TRS-003; TRS-050 | Inspect placement, anchoring, clearances, floor loading vs vendor drawings and DQ. | Installed per DQ drawings; anchoring and clearances meet spec; floor loading acceptable. | Site layout drawing; structural calc certs. | Once | Engineering |
| IQ-04 | Utilities connections — TRS-014; TRS-012; TRS-066; TRS-051 | Verify steam, condensate, electrical, vacuum, compressed air connections per spec; visual leak check. | Connections per spec; no visible leaks; correct voltage/phase; isolation valves present. | Visual; calibrated clamp meter; once | Facilities |  |
| IQ-05 | Steam inlet verification — TRS-014; TRS-013; TRS-103 | Measure steam inlet pressure and temperature at idle and verify sampling port accessibility. | Steam inlet 3–6 bar at idle; sampling ports accessible; separator/trap installed per TRS. | Calibrated pressure gauge; thermocouple; once | Facilities |  |
| IQ-06 | Drain and condensate routing — TRS-051; TRS-069 | Inspect drain slope, trap, condensate return, and condensate treatment connections. | Drain flows freely; traps installed; no backflow; condensate routing per DQ. | Visual inspection; once | Facilities |  |
| IQ-07 | Safety devices installed — TRS-021; TRS-022; TRS-020 | Verify door interlock, relief valve, emergency stop, horn/strobe, and alarm wiring installed and functional. | Devices installed, wired, and functional; alarms mapped to HMI. | Functional test; calibrated pressure gauge for relief valve; once | Engineering; Safety |  |
| IQ-08 | Instrumentation installed — TRS-008; TRS-009; TRS-010; TRS-011 | Confirm temperature, pressure, vacuum sensors, redundant probes, and mounting ports at DQ locations. | All sensors present at specified locations; redundant sensors installed; wiring to controller verified. | Visual; sensor certs attached; once | Engineering |  |
| IQ-09 | Network, software, and firmware baseline — TRS-015; TRS-016; TRS-023; TRS-024 | Verify controller version, network config, audit DB, backups scheduled, and firmware image repository. | Controller version matches DQ; NTP and backup configured; audit DB present; firmware linked to V&V. | System printout; backup config export; once | IT; QA |  |
| IQ-10 | Calibration certificates — TRS-035; TRS-055 | Collect calibration certificates for all validation instruments and critical sensors including measurement uncertainty. | All instruments have current certs traceable to accredited lab; uncertainty fields present. | Calibration certificates; once | Validation Lead; QA |  |
| IQ-11 | Spare parts and consumables — TRS-033; TRS-080; TRS-082 | Verify critical spares list, consumable master data, and minimum on-hand quantities. | Critical spares on site or PO placed; consumable qualification records attached. | PO/inventory list; once | Procurement |  |
| IQ-12 | SOPs, training, and gating — TRS-042; TRS-041; TRS-089 | Confirm SOPs for operation, cleaning, maintenance, BI/CI handling, and validation exist; training matrix configured and gating enforced. | SOPs approved; training plan in place; system gating configured to block untrained users. | SOP documents; training matrix; once | QA; Training |  |
| IQ-13 | Condensate/steam system components — TRS-013; TRS-128; TRS-095 | Inspect steam separator, traps, steam trap/separator maintenance access, and sampling ports. | Separator and traps installed per vendor spec; sampling ports accessible; maintenance access provided. | Visual; once | Facilities; Maintenance |  |
| IQ-14 | Alarm and HMI verification — TRS-020; TRS-052 | Verify HMI alarm pages, horn/strobe, priority mapping, and SOP link integration. | HMI alarm pages present; audible/visual outputs functional; SOP links accessible from alarm pages. | Functional alarm simulation; once | Engineering; IT |  |
| IQ-15 | Backup and archive configuration — TRS-039; TRS-040; TRS-054 | Verify backup schedule, archive repository, export capability, and checksum generation configured. | Backups scheduled; archive accessible; export and checksum functions verified. | Backup/restore test printout; once | IT |  |
| IQ-16 | BD/air removal program and vacuum hardware — TRS-006; TRS-012; TRS-057 | Verify BD program present, vacuum pump and valves installed, and leak test ports accessible. | BD program available; vacuum hardware installed; leak test ports accessible. | Visual; software module printout; once | Engineering |  |
| IQ-17 | Mounting fixtures and mapping setup — TRS-036; TRS-030; TRS-035 | Confirm data logger mounting fixtures, mapping coordinates, and mounting SOPs are available. | Mounting fixtures installed; mapping coordinates documented; photographic SOP present. | Visual; photos of mounting fixtures; once | Validation Lead; Engineering |  |
| IQ-18 | Labeling and traceability hardware — TRS-079; TRS-045 | Verify label printer integration, barcode formats, and scanner connectivity to load record schema. | Printer drivers installed; sample labels print correctly; barcode scan populates load record fields. | Print/scan test; once | IT; Production |  |
| IQ-19 | Environmental and HVAC interfaces — TRS-068; TRS-049 | Verify HVAC connections, environmental sensor installation for storage/handling areas, and interface wiring. | HVAC interfaces present; environmental sensors installed and reporting to system. | Visual; sensor readout; once | Facilities |  |
| IQ-20 | Ready for OQ sign-off — TRS-031; TRS-032 | QA and Engineering review IQ evidence and sign readiness for OQ. | QA and Engineering sign off that IQ items passed or CAPA in place for acceptable deviations. | IQ checklist summary; once | QA; Engineering |  |

Deviation Handling during IQ:

Record any nonconformance on the deviation form. Assess impact on sterility, safety, and validated configuration. If safety or sterility risk exists, stop work, initiate containment, and open CAPA. Rework and retest affected IQ items; attach retest evidence.

IQ Signatures

Prepared by: Validation Lead — Date: ________

Reviewed by: Engineering Manager — Date: ________

Approved by: QA Manager — Date: ________

Operational Qualification (OQ) Protocol

Purpose:

Demonstrate the sterilizer functions per the Functional Requirement Specifications (FRS) across the operational envelope and verify controls, alarms, software, and functional workflows behave as required.

Prerequisites:

IQ complete and signed; calibrated validation instruments on site with traceable certificates; FAT/SAT evidence attached to DQ/IQ; SOPs available; microbiology controls (BI/CI) and lots ready; trained personnel scheduled; data logger mounting fixtures and mapping coordinates defined.

Execution Notes:

Define mapping coordinates and mounting method before testing. Use calibrated data loggers and reference probes with traceable certificates. Run empty chamber mapping first, then loaded worst-case mapping per load templates. Capture raw data files, export checksums, and attach calibration certificates and BI/CI lot certificates. Record deviations immediately and route to CAPA if they affect sterility, safety, or validated configuration. Include statistical summaries (mean, SD, Cp/Cpk where applicable) in OQ report.

OQ Protocol Table

| Test ID | Requirement (FRS mapping) | Test Method | Acceptance Criterion | Instrument Calibration | Frequency | Responsible |
| --- | --- | --- | --- | --- | --- | --- |
| OQ-001 | FRS-001: Validated sterilization workflows for each load type | Execute each workflow with representative empty and loaded runs; record cycle logs and BI/CI placement | Workflows execute without error; cycle logs include cycle ID, load map, BI/CI placements; no unauthorized recipe changes | Calibrated data loggers; thermocouples; pressure gauges | Once per OQ; repeat after changes | Validation Lead; Production |
| OQ-002 | FRS-002: Packaging format processing and drying | Run cycles for single/double wrap, rigid containers, peel pouches; inspect seals and CI response; measure residual moisture | CI change as expected; no seal failures; residual moisture ≤3% or CI dry-pack pass | Calibrated moisture balance; thermocouples | Once per packaging type | Validation Lead; QA |
| OQ-003 | FRS-003: Cycle performance at max gross load (50 kg) | Run worst-case 50 kg loaded cycle with mapping and BIs at worst-case positions | BI negative; mapped T/P/vacuum within acceptance; controller logs within limits | Calibrated data loggers; calibrated scale | Once | Validation Lead; Engineering |
| OQ-004 | FRS-004: Load template enforcement and rack compatibility | Select each load template and run mapping; verify HMI records template ID and coordinates | System records correct template; physical placement matches template; no template errors | Visual; data logger certs | Once per template | Production; IT |
| OQ-005 | FRS-005: Pre-vacuum recipe control (134°C/3 min; 132°C/4 min) | Execute each recipe 3 times; verify setpoint attainment, plateau, and F₀ calculation | Setpoint within ±0.5°C; plateau time met; F₀ meets lethality target; repeatable across runs | Calibrated reference thermocouple; pressure transducer | 3 runs each | Validation Lead; Engineering |
| OQ-006 | FRS-006: Bowie-Dick and dynamic air removal diagnostics | Run BD program per standard (3 consecutive) and record vacuum profile | BD passes per acceptance; vacuum profile within defined limits | BD test packs; vacuum gauge calibration | 3 consecutive runs | Validation Lead; Engineering |
| OQ-007 | FRS-007: RBAC and parameter limit enforcement | Attempt recipe edits as unauthorized and authorized users; verify audit trail entries | Unauthorized edits blocked; authorized edits logged with reason and user ID | System audit log printout | Once | IT; QA |
| OQ-008 | FRS-008: Setpoint accuracy ±0.5°C | Hold plateau and compare controller to calibrated reference probe at multiple points | Controller vs reference ≤ ±0.5°C at plateau for 3 repeats | Calibrated thermocouple (traceable) | 3 repeats | Validation Lead; Engineering |
| OQ-009 | FRS-009: Temperature uniformity ±2.0°C | Empty and loaded mapping with data loggers at defined coordinates; 3 runs | All mapped points within ±2.0°C of setpoint for empty and loaded conditions | Calibrated data loggers | 3 runs empty; 3 runs loaded | Validation Lead |
| OQ-010 | FRS-010: Redundant sensor discrepancy detection | Simulate sensor failure/discrepancy and verify alarm and fallback behavior | Discrepancy alarm generated; system follows fallback logic and logs event | Calibrated thermocouples; system alarm log | Once per sensor set | Engineering; QA |
| OQ-011 | FRS-011: Pressure/vacuum accuracy ±2% FS | Compare controller readings to calibrated reference gauges across range | Readings within ±2% FS at low, mid, high points | Calibrated pressure/vacuum gauges | 3 points across range | Engineering |
| OQ-012 | FRS-012: Vacuum capability ≤50 mbar and leak rate ≤5 mbar/min | Evacuate chamber to target and perform vacuum hold/leak test | Achieve ≤50 mbar and leak rate ≤5 mbar/min | Calibrated vacuum gauge | Once; repeat after maintenance | Engineering |
| OQ-013 | FRS-013: Steam quality sampling (dryness ≥0.95; NCG ≤3.5%) | Collect steam samples during cycle and send to lab for dryness/NCG analysis | Lab results meet dryness and NCG limits | Sampling kit calibration; lab certificate | 1 sample per OQ cycle | Facilities; QA |
| OQ-014 | FRS-014: Steam inlet pressure monitoring (3–6 bar) | Monitor inlet pressure during idle and cycle; simulate excursions | Pressure maintained 3–6 bar; system alarms and blocks outside range | Calibrated pressure gauge | Continuous during OQ runs | Facilities |
| OQ-015 | FRS-015: Data logging, export, and archival verification | Export cycle logs, verify checksum, and restore from archive | Exported files readable; checksum matches; restore successful | Data logger certs; IT backup logs | Once per OQ | IT; QA |
| OQ-016 | FRS-016: Audit trail immutability and metadata | Perform critical actions and verify audit entries include user ID, timestamp, reason and are immutable | Audit entries present and cannot be altered; searchable | System audit DB printout | Once | IT; QA |
| OQ-017 | FRS-017: RBAC and password policy enforcement tests | Create test users, enforce password complexity, test lockout and provisioning | Password policy enforced; lockout occurs after defined attempts; provisioning logged | System user management export | Once | IT |
| OQ-018 | FRS-018: Electronic signature linkage test | Apply e-signature to test record and verify metadata and linkage | E-signature recorded with user ID, reason, timestamp and linked to record | System e-signature module; audit log | Once (if applicable) | IT; QA |
| OQ-019 | FRS-019: Secure export and checksum/digital signature verification | Export sample cycle file and verify checksum/digital signature integrity | Exported file checksum/digital signature verifies and matches archive | Export tool; checksum utility | Once | IT |
| OQ-020 | FRS-020: Alarm classification, HMI guidance, and escalation | Simulate alarms of different priorities and verify HMI guidance, logs, and escalation notifications | Alarms display correct priority, HMI shows operator guidance, escalation notifications sent | Alarm simulation tool; HMI logs | Multiple simulated events | Engineering; QA |
| OQ-021 | FRS-021: Door interlock functional test | Simulate unsafe conditions and verify door remains locked and attempts logged | Door cannot open until safe; attempt logged with user/timestamp | Door position sensor; interlock test rig | Once | Engineering; Safety |
| OQ-022 | FRS-022: Emergency stop and power loss safe-state test | Activate E-stop and simulate power loss; verify safe state and data preservation | E-stop triggers safe state; UPS preserves controller data; event logged | UPS test; data integrity check | Once | Engineering; IT |
| OQ-023 | FRS-023: Firmware version verification and V&V linkage | Verify installed firmware matches approved V&V package and record version in logs | Firmware version matches V&V; version recorded in system | System printout; V&V docs | Once | IT; QA |
| OQ-024 | FRS-024: Software function tests and Part 11 controls | Execute software test scripts per V&V and Part 11 checklist | All software functions pass V&V tests; Part 11 controls operate as required | Software test harness; audit log | Once | IT; QA |
| OQ-025 | FRS-025: SPC and reproducibility tests | Run repeated cycles (n≥5) and generate SPC charts for key parameters | Parameters within control limits; no out-of-control signals; SD acceptable | Calibrated data loggers; SPC software | 5+ consecutive cycles | Validation Lead; QA |
| OQ-026 | FRS-026: BI placement plan execution and recording | Place BIs per plan and run cycles; verify BI location recorded in load record | BI locations recorded correctly and linked to cycle ID | BI lot certs; incubator | Per BI plan | Microbiology; Production |
| OQ-027 | FRS-027: CI placement and result capture | Place CIs in critical locations and run cycles; verify CI result capture | CI results logged with pass/fail flags and linked to load record | CI lot certs | Per CI plan | Production; QA |
| OQ-028 | FRS-028: Drying control and residual moisture verification | Run drying cycles and perform gravimetric moisture tests on sample packs | Residual moisture ≤3% or CI dry-pack pass across repeated runs | Calibrated moisture balance | Multiple drying runs | Validation Lead; Production |
| OQ-029 | FRS-029: Drying reproducibility analysis | Execute repeated drying cycles and analyze residual moisture SD and reproducibility | Residual moisture SD within defined limits and reproducible across runs | Moisture balance; statistical tool | 5+ runs | Validation Lead |
| OQ-030 | FRS-030: Worst-case load configuration mapping | Execute worst-case load patterns with mapping and BIs at worst-case positions | Mapping shows acceptable T/P/vacuum; BIs negative at worst-case positions | Calibrated data loggers; BI lot certs | Once per load pattern | Validation Lead; Production |
| OQ-031 | FRS-031: PQ workflow verification (pre-PQ gating) | Execute PQ workflow module in test mode to enforce three consecutive loads and gating | PQ workflow enforces sequence and prevents release until all passes | System PQ module printout | Once | Validation Lead; QA |
| OQ-032 | FRS-032: Requalification trigger simulation | Create change control entry and verify requalification trigger and scheduling | Change control triggers requalification workflow and schedules tasks | Change control module printout | Once | QA; Validation Lead |
| OQ-033 | FRS-033: Spare parts alert test | Simulate part depletion and verify reorder alert generation | Reorder alert generated when stock below min; notification sent | Spare parts DB; inventory system | Once | Procurement |
| OQ-034 | FRS-034: PM task execution and parameter protection | Execute PM task and verify record creation and that validated parameters cannot be changed without change control | PM record created; parameter change blocked without change control | CMMS printout | Once per PM test | Maintenance; QA |
| OQ-035 | FRS-035: Calibration record and uncertainty verification | Review calibration certs and verify uncertainty fields used in acceptance calculations | Calibration certs present and uncertainty applied in acceptance logic | Calibration certificates | Once | Validation Lead; QA |
| OQ-036 | FRS-036: Data logger mounting reproducibility test | Mount data loggers per SOP and perform mapping; verify photographic evidence and repeatability | Photos attached; mapping repeatable across runs | Calibrated data loggers; camera | 3 mapping runs | Validation Lead |
| OQ-037 | FRS-037: Utility failure detection and notification test | Simulate steam loss/low pressure and verify alarms, notifications, and safe shutdown | Alarms triggered; notifications sent per SOP; safe shutdown if required | Utility sensors; alarm logs | Simulated events | Facilities; QA |
| OQ-038 | FRS-038: Vulnerability scan and IT security checklist verification | Run vulnerability scan and verify remediation of critical findings | No critical vulnerabilities; IT checklist items met | Vulnerability scan report | Once | IT; Security |
| OQ-039 | FRS-039: Backup and restore verification | Perform backup and restore of controller data and verify integrity and RTO/RPO | Restore successful within RTO/RPO and checksum matches | Backup/restore logs; checksum tool | Once | IT; QA |
| OQ-040 | FRS-040: Archive retrieval test | Generate archive index and retrieve sample validation package within SLA | Retrieval returns complete dataset within defined SLA | Archive retrieval test report | Once | QA; IT |
| OQ-041 | FRS-041: Training gating enforcement test | Attempt cycle start with untrained user (expect block) and with trained user (allow) | Untrained user blocked; trained user allowed and action logged | Training matrix export; system logs | Once | Training; QA |
| OQ-042 | FRS-042: HMI SOP confirmation enforcement | Execute critical step and verify HMI displays SOP link and requires confirmation | HMI prompts SOP confirmation and logs confirmation before proceeding | HMI logs; SOP link verification | Once per critical step | QA; IT |
| OQ-043 | FRS-043: BI/CI lot control enforcement | Attempt to use expired BI/CI lot (expect block) and use valid lot (accept) | System blocks expired lots and accepts valid lots with certificate attachment | BI/CI lot certs; system logs | Once | Microbiology; QA |
| OQ-044 | FRS-044: BI positive/negative control incubation verification | Incubate positive and negative controls and verify expected outcomes | Positive control shows growth; negative control shows no growth | Incubator calibration cert | Per PQ run | Microbiology |
| OQ-045 | FRS-045: Load traceability scan and record verification | Scan device lot barcodes into load record and verify fields populated and retrievable | Load record contains device lot, operator, cycle ID, BI/CI results and is retrievable | Barcode scanner calibration; system logs | Once per load type | Production; QA |
| OQ-046 | FRS-046: FAT/SAT deliverable linkage verification | Review FAT/SAT documents linked to DQ/IQ and verify accessibility | FAT/SAT documents accessible and linked to equipment records | Document repository printout | Once | Procurement; QA |
| OQ-047 | FRS-047: Material certificate verification and surface finish inspection | Review material certificates and inspect chamber surfaces for specified finish | Certificates present and surface finish meets spec | Material certs; visual inspection | Once | Procurement; QA |
| OQ-048 | FRS-048: Cleaning SOP execution and pre-PQ verification | Execute cleaning SOP and verify cleaning log entry and visual cleanliness | Cleaning log completed and visual inspection acceptable prior to PQ | Cleaning log; visual checklist | Once per cleaning event | Production; QA |
| OQ-049 | FRS-049: Ambient T/H monitoring during cycles | Record ambient temperature/humidity during OQ cycles and verify no adverse effect | Ambient T/H within defined range or mitigations documented; no adverse effect on cycle | Calibrated ambient sensors | During OQ runs | Facilities; QA |
| OQ-050 | FRS-050: Installation anchoring and clearance verification | Confirm anchoring and clearances remain per IQ during operation | No movement; clearances maintained during operation | Visual inspection; once | Engineering |  |
| OQ-051 | FRS-051: Condensate routing functional test | Run cycles and observe condensate routing and trap function under load | Condensate routed correctly; no backflow; traps functioning | Visual; condensate flow logs | During OQ runs | Facilities |
| OQ-052 | FRS-052: Audible/visual alarm and escalation simulation | Simulate critical alarm and verify horn/strobe, HMI, and escalation notifications | Audible/visual alarms activate; escalation notifications sent per SOP | Alarm simulation tool; HMI logs | Multiple simulated events | Engineering; QA |
| OQ-053 | FRS-053: Partial requalification script execution | Simulate minor repair and execute partial requalification tests per template | Partial requalification tests pass and sign-offs recorded | Test scripts; calibrated instruments | After minor repair | Validation Lead; QA |
| OQ-054 | FRS-054: Raw data export and checksum verification | Export raw data files and verify metadata and checksum before archival | Raw files exported with metadata and checksum; archive ingestion verified | Export tool; checksum utility | Once per OQ run | IT; Validation Lead |
| OQ-055 | FRS-055: Measurement uncertainty application in acceptance | Apply instrument uncertainty to acceptance calculations for mapping and pressure tests | Acceptance decisions documented with uncertainty propagation | Calibration certs with uncertainty | During OQ analysis | Validation Lead; QA |
| OQ-056 | FRS-056: Lumen penetration tests with BI at lumen tips | Execute lumen tests for each lumen dimension and incubate BIs | BIs negative at lumen tips for validated lumen dimensions | BI lot certs; lumen fixtures | Per lumen matrix | Microbiology; Validation Lead |
| OQ-057 | FRS-057: BD/equivalent test scaled to chamber size | Run BD/equivalent test and record pass/fail per defined criteria | BD passes per acceptance criteria for chamber size | BD packs; vacuum gauge | 3 consecutive runs | Validation Lead |
| OQ-058 | FRS-058: Change control trigger verification | Create change control entry and verify requalification trigger and notifications | Change control evaluates impact and triggers requalification workflow when criteria met | Change control module printout | Once | QA |
| OQ-059 | FRS-059: Deviation/CAPA workflow test | Create deviation and route through CAPA to closure with evidence | Deviation routed, CAPA actions recorded, verification evidence attached and closed | CAPA module printout | Once | QA |
| OQ-060 | FRS-060: Final validation report generation test | Generate draft validation report including traceability matrix and attachments | Report contains traceability matrix, raw data attachments, BI/CI results, deviations, CAPA, and signature fields | Report generator output | Once | Validation Lead; QA |
| OQ-061 | FRS-061: Material compatibility test execution | Run pre/post cycle material tests (visual, mechanical, chemical) per plan | No unacceptable degradation beyond defined limits; results recorded | Test equipment calibration certs | Per material plan | R&D; Materials Eng; QA |
| OQ-062 | FRS-062: Packaging matrix and seal peel tests | Execute packaging tests and seal peel tests per matrix | Seal integrity maintained; CI/BI pass; peel strength within limits | Peel tester calibration; CI/BI certs | Per packaging matrix | Packaging Eng; QA |
| OQ-063 | FRS-063: Lumen/geometry matrix execution | Execute lumen and nested geometry tests per matrix and record BI results | BI negative for validated geometries and nested configs | Lumen fixtures; BI certs | Per matrix | Microbiology; Validation Lead |
| OQ-064 | FRS-064: Assembled/disassembled state validation and functional checks | Run cycles for both states and perform functional tests post-cycle | Both states meet BI/CI and functional acceptance criteria | Functional test fixtures; BI certs | Per device state | Production; Validation Lead |
| OQ-065 | FRS-065: Stability study baseline testing | Execute initial stability tests and record baseline sterility and packaging integrity | Baseline sterility and packaging integrity meet acceptance for stability plan | Stability test equipment | Once (baseline) | Regulatory Affairs; QA |
| OQ-066 | FRS-066: Water/condensate sampling and alarm verification | Sample feedwater/condensate and verify lab results and alarm behavior on excursions | Lab results within limits; alarms generated for excursions | Sampling kit; lab certs | Per OQ cycle | Facilities; QA |
| OQ-067 | FRS-067: Altitude correction verification | Run cycles with altitude correction applied and verify lethality metrics | Adjusted parameters documented and BI acceptance met at site altitude | Controller printout; BI certs | Once | Validation Lead; Engineering |
| OQ-068 | FRS-068: HVAC/environmental interface verification for storage | Verify environmental data linkage to storage lots and alarm behavior | Environmental data linked to lot records and alerts generated on excursions | Environmental sensors; system logs | Once | Facilities; QA |
| OQ-069 | FRS-069: Condensate/exhaust compliance sampling | Collect condensate/exhaust samples and attach lab reports for compliance check | Sample results compared to permit limits and documented | Sampling kit; lab certs | Per OQ plan | Facilities; EHS |
| OQ-070 | FRS-070: Patch management predeployment test | Test patch in non-production environment and verify rollback plan and change control | Patch tested and rollback verified; change control approval recorded | Test environment logs | Per patch | IT; QA |
| OQ-071 | FRS-071: NTP/time sync verification | Verify NTP sync and measure timestamp drift over defined interval | Timestamps synchronized and drift within tolerance | NTP logs; time server printout | 24-hour drift check | IT |
| OQ-072 | FRS-072: Encryption and secure transfer verification | Test TLS and disk encryption and verify secure transfer logs | Encryption active and secure transfer logs recorded and verifiable | Security config printout | Once | IT; Security |
| OQ-073 | FRS-073: MES/LIMS integration test | Execute API transfers and reconciliation tests for sample cycles | Data mapping correct and reconciliation passes; no data loss | API test logs; reconciliation report | Integration test | IT; Production |
| OQ-074 | FRS-074: DR restore test | Perform restore test and measure recovery time vs RTO/RPO | Restore completes within RTO/RPO and data integrity verified | Backup/restore logs | Once | IT |
| OQ-075 | FRS-075: Lot release interlock test with BI/CI gating | Simulate BI/CI fail and pass and verify hold/release logic and MES handoff | Hold prevents release on fail; release occurs automatically on pass and MES receives release | System release logs; MES reconciliation | Per release workflow | QA; IT |
| OQ-076 | FRS-076: Sampling plan execution and acceptance calculation | Execute statistical sampling plan and verify acceptance calculations | Sampling results recorded and acceptance decision matches plan | Statistical tool; sampling records | Per sampling plan | QA; Biostatistics |
| OQ-077 | FRS-077: Indicator selection and lot blocking test | Attempt to select non-conforming indicator (expect block) and select conforming indicator | System blocks non-conforming lots and accepts conforming lots with certs | Indicator master data; system logs | Once | Microbiology; QA |
| OQ-078 | FRS-078: Environmental monitoring trend generation | Simulate environmental data and generate trend reports | Trend reports generated per cadence and include alerts for excursions | Environmental data logs; trend tool | Once | QA; Facilities |
| OQ-079 | FRS-079: Label print and barcode verification | Print sample labels and scan to populate load/pack records; verify accuracy | Labels print correctly and barcode scan populates load record fields accurately | Label printer calibration; barcode scanner | Once per label format | Production; IT |
| OQ-080 | FRS-080: Spare parts reorder alert test | Deplete part and verify reorder alert and notification routing | Reorder alert generated and notification sent to procurement | Spare parts DB; notification logs | Once | Procurement |
| OQ-081 | FRS-081: Maintenance trigger and change control linkage test | Execute maintenance action that should trigger requalification and verify change control initiation | Change control initiated and requalification trigger recorded | Maintenance record; change control module | Once | Maintenance; QA |
| OQ-082 | FRS-082: Consumable qualification enforcement test | Attempt to use unqualified consumable (expect block) and use qualified consumable (accept) | System blocks unqualified consumables and accepts qualified ones with certs | Consumable master data; system logs | Once per consumable type | Procurement; QA |
| OQ-083 | FRS-083: Decommissioning module test (mock) | Execute mock decommissioning and verify archive export and retention metadata | Decommissioning checklist executed and records archived with retention metadata | Archive export logs | Mock test | QA |
| OQ-084 | FRS-084: Reportable event classification and notification test | Simulate reportable event and verify notification workflow and timelines | Notification workflow triggered and timelines met per SOP | Event log; notification timestamps | Simulated event | Regulatory Affairs; QA |
| OQ-085 | FRS-085: Emergency drill execution and logging | Conduct emergency drill and verify drill log, notifications, and corrective actions | Drill log captured with outcomes and corrective actions recorded | Drill log; notification records | Per drill schedule | QA; EHS |
| OQ-086 | FRS-086: Audit pack generation and completeness test | Generate audit pack and verify inclusion of required documents and retrieval time | Audit pack complete per template and retrieval within SLA | Audit pack generator output | Once | QA |
| OQ-087 | FRS-087: Permit repository and expiration alert test | Upload permit and simulate expiration alert; verify workflow | Expiration alert generated and workflow initiated | Permit repository logs | Once | Regulatory Affairs |
| OQ-088 | FRS-088: Consumable/spare substitution change control test | Submit substitution change and verify impact assessment and verification test scheduling | Impact assessment recorded and verification tests scheduled and executed | Change control record | Once per substitution | Procurement; QA |
| OQ-089 | FRS-089: Recurring training scheduling and gating test | Schedule training and competency requalification and verify access gating | Training scheduled; competency recorded; access gated until completion | Training system logs | Once | Training; QA |
| OQ-090 | FRS-090: Trend analytics generation and management review export | Generate trend report for BI/CI failures and deviations and export for management review | Trend report generated and management review export available with action items | Trend tool output | Once | QA |
| OQ-091 | FRS-091: Emergency cycle limited validation workflow test | Execute emergency cycle limited validation and verify risk assessment and approvals | Emergency cycle limited tests executed and approvals recorded per workflow | Emergency cycle test scripts | Test event | Validation Lead; QA |
| OQ-092 | FRS-092: Third-party service provider record verification | Review sample service report and verify provider qualification and sign-off fields | Service report attached and provider qualification on file | Service report; supplier records | Once | Maintenance; Procurement |
| OQ-093 | FRS-093: Accessory compatibility test in load mapping | Include accessory in load and run mapping and BI/CI tests | Accessories do not adversely affect mapping or BI/CI results | Mapping data loggers; BI certs | Per accessory | Production; Validation Lead |
| OQ-094 | FRS-094: Nonconforming product workflow test | Create nonconforming event and verify quarantine, investigation, and disposition workflow | Quarantine flag set and disposition workflow executed with evidence | Nonconforming module logs | Simulated event | QA |
| OQ-095 | FRS-095: Trap/separator maintenance and post-maintenance steam quality test | Perform maintenance and sample steam quality post-maintenance | Post-maintenance steam sample meets quality limits | Sampling kit; lab certs | After maintenance | Facilities; Maintenance |
| OQ-096 | FRS-096: MSA/Gage R&R execution and reporting | Conduct Gage R&R for data loggers/thermocouples and analyze results | MSA results meet capability criteria and report stored | MSA templates; calibrated instruments | Once | Validation Lead; QA |
| OQ-097 | FRS-097: VMP annual review workflow test | Execute VMP review and verify revision capture and approvals | VMP review minutes captured and revisions recorded with approvals | Management review record | Annual | QA |
| OQ-098 | FRS-098: Supplier QA agreement and deliverable verification | Review supplier QA agreement and verify FAT/SAT and spares deliverables present | Supplier deliverables present and accepted per checklist | Supplier agreement; FAT/SAT docs | Once | Procurement; QA |
| OQ-099 | FRS-099: Environmental extreme condition OQ tests | Execute OQ tests under simulated ambient extremes or verify mitigations | Cycle performance acceptable under extremes or mitigations documented and effective | Environmental test harness | Per site needs | Facilities; Validation Lead |
| OQ-100 | FRS-100: Uncertainty propagation verification in acceptance | Perform acceptance calculations including measurement uncertainty and document decisions | Acceptance decisions include uncertainty and are justified in report | Calibration certs with uncertainty | During OQ analysis | Validation Lead; QA |
| OQ-101 | FRS-101: Pre-cleaning bioburden sampling verification | Perform incoming bioburden sampling and verify results against limits | Bioburden results within defined limits and pre-cleaning SOPs followed | Microbiology lab certs | Per sampling plan | Production; Microbiology |
| OQ-102 | FRS-102: Approved detergents compatibility test | Test residue and compatibility after cleaning with approved detergents | Residue within acceptable limits and no adverse effects observed | Chemical analysis certs | Per detergent | R&D; QA |
| OQ-103 | FRS-103: Steam source validation for each source | Run cycles using each steam source and sample steam quality and cycle performance | Cycle performance and steam quality meet URS limits for each source | Sampling kit; BI certs | Per steam source | Facilities; Validation Lead |
| OQ-104 | FRS-104: UPS failover and data preservation test | Simulate power interruption and verify UPS failover and data preservation | UPS supports safe shutdown and data preserved; logs show event | UPS test logs | Once | IT; Facilities |
| OQ-105 | FRS-105: EMC/grounding site verification | Perform site EMC/grounding checks and verify no controller anomalies | EMC/grounding checks pass and no EMI anomalies observed | EMC test report; grounding check tools | Once | Engineering; IT |
| OQ-106 | FRS-106: Noise/vibration measurement during operation | Measure noise and vibration and compare to limits | Noise/vibration within occupational and equipment limits | Calibrated sound/vibration meters | Once | Facilities; EHS |
| OQ-107 | FRS-107: Seasonal ambient mitigation verification | Execute seasonal tests or verify mitigations and document results | Performance acceptable or mitigations effective and documented | Environmental test harness | Per site | Facilities; Validation Lead |
| OQ-108 | FRS-108: MSA/Gage R&R repeatability/reproducibility verification | Conduct MSA/Gage R&R and verify repeatability/reproducibility metrics | MSA meets predefined capability criteria | MSA templates; calibrated instruments | Once | Validation Lead; QA |
| OQ-109 | FRS-109: Sensor drift detection and recalibration trigger test | Simulate sensor drift and verify detection and recalibration workflow trigger | Drift detected and recalibration workflow initiated and recorded | Drift simulation tool; calibration certs | Once per sensor type | Engineering; QA |
| OQ-110 | FRS-110: Firmware rollback test in non-production | Perform controlled rollback test and verify data integrity and recovery | Rollback successful without data loss in test environment | Firmware repo; test logs | Test event | IT; Supplier |
| OQ-111 | FRS-111: Remote vendor access control and logging test | Simulate vendor remote session and verify approval, time limits, and session logs | Remote session required approval, time limited, and logs stored | Remote access gateway logs | Simulated session | IT; QA |
| OQ-112 | FRS-112: Mobile device access and audit logging test | Perform mobile login and actions and verify audit trail entries | Mobile actions authenticated and logged in audit trail | Mobile device management logs | Once | IT; QA |
| OQ-113 | FRS-113: MES/LIMS release interlock integration test | Simulate BI/CI pass/fail and verify MES/LIMS interlock and reconciliation | MES/LIMS receives release only after BI/CI pass and approvals; reconciliation matches | API logs; reconciliation report | Integration test | IT; QA |
| OQ-114 | FRS-114: Retained sample retrieval test | Create retention record and retrieve sample to verify retrieval logging | Retrieval logged and sample integrity verified | Retention system logs | Once | QA |
| OQ-115 | FRS-115: Post-maintenance verification script execution | Perform major maintenance and execute post-maintenance verification tests | Verification tests pass and production hold cleared only after sign-off | Verification scripts; calibrated instruments | After major maintenance | Maintenance; Validation Lead |
| OQ-116 | FRS-116: Obsolescence register and alternative part verification | Review obsolescence entries and verify alternative part approval records | Alternative parts approved with impact assessment and verification evidence | Obsolescence register | Once | Procurement; Engineering |
| OQ-117 | FRS-117: Orientation/stacking validation tests | Run orientation and stacking patterns with BI/CI placement and mapping | BI/CI negative and mapping within acceptance for validated orientations | Calibrated data loggers; BI certs | Per orientation pattern | Production; Validation Lead |
| OQ-118 | FRS-118: Mixed-load validation tests | Execute mixed-load patterns and verify BI/CI and mapping results across device types | BI/CI negative and mapping acceptable for mixed loads | BI certs; data loggers | Per mixed-load pattern | Production; Validation Lead |
| OQ-119 | FRS-119: Implant workflow and enhanced documentation verification | Execute implant workflow including enhanced traceability and regulatory checklist | Enhanced documentation complete and regulatory checklist signed | Workflow logs; regulatory checklist | Per implant load | Regulatory Affairs; QA |
| OQ-120 | FRS-120: Endotoxin test workflow execution (if applicable) | Perform endotoxin testing on representative lots and record results | Endotoxin results within acceptance where required | Endotoxin test kit calibration | Per requirement | QA; Microbiology |
| OQ-121 | FRS-121: Post-move IQ/OQ verification after relocation | Execute post-move IQ/OQ scripts and vibration/transport inspection | Post-move IQ/OQ pass and transport inspection records attached | Vibration meter; IQ/OQ scripts | After relocation | Validation Lead; Engineering |
| OQ-122 | FRS-122: Cross-site comparability assessment | Compare OQ data to reference site and document comparability | Site data comparable or deviations justified and mitigations documented | OQ data exports | Once per site | Corporate Validation; QA |
| OQ-123 | FRS-123: Supplier audit evidence verification | Review supplier audit reports and corrective action evidence for critical suppliers | Supplier audit findings closed or mitigations documented | Supplier audit reports | Once | Procurement; QA |
| OQ-124 | FRS-124: Accessory cleaning validation and PQ inclusion verification | Execute accessory cleaning validation and include accessories in OQ mapping | Accessories pass cleaning validation and included in mapping with acceptable BI/CI | Cleaning validation records; BI certs | Per accessory | Production; Validation Lead |
| OQ-125 | FRS-125: Raw data provenance and chain of custody verification | Export raw files and verify metadata, checksum, and chain of custody entries | Raw files include metadata and checksum; chain of custody complete | Export tool; checksum utility | Once per OQ run | IT; QA |
| OQ-126 | FRS-126: Redaction workflow test for external record sharing | Redact sample record and verify retained validation evidence and approval log | Redacted record retains required evidence and redaction log shows approvals | Redaction tool logs | Once | QA |
| OQ-127 | FRS-127: Water hardness test and mitigation linkage verification | Perform water hardness tests and verify mitigation controls applied when outside limits | Hardness results recorded and mitigation controls applied and documented | Hardness test kit | Per sampling plan | Facilities; QA |
| OQ-128 | FRS-128 | Perform trap/separator maintenance and post-maintenance steam sampling | Post-maintenance steam sample meets quality limits and records attached | Sampling kit; maintenance records | After maintenance | Facilities; Maintenance |
| OQ-129 | FRS-129: Humidity variation drying tests | Run drying tests under varied humidity conditions and verify residual moisture | Residual moisture within acceptance across defined humidity range or mitigations documented | Calibrated humidity sensors; moisture balance | Per test plan | Validation Lead; Facilities |
| OQ-130 | FRS-130: Embedded electronics functional pre/post tests | Execute functional tests on electronics before and after sterilization per vendor guidance | No functional degradation; results meet vendor acceptance | Functional test fixtures; calibration certs | Per device | R&D; QA |
| OQ-131 | FRS-131: CI class interpretation verification | Run CI class tests and verify interpretation recorded per SOP in PQ records | CI interpretation matches SOP and recorded in system | CI lot certs | Per CI class test | Microbiology; QA |
| OQ-132 | FRS-132: Alternative cycle recipe OQ tests | Execute OQ/PQ for alternative cycles and perform material compatibility tests | Alternative cycle meets BI/CI and material compatibility acceptance | BI certs; material test equipment | Per alternative cycle | Validation Lead; R&D |
| OQ-133 | FRS-133: Timestamp integrity and DST handling verification | Verify NTP sync and timestamp handling including DST/leap seconds | Timestamps synchronized and DST/leap second handling correct | NTP logs | 24-hour check | IT |
| OQ-134 | FRS-134: Seal replacement and leak test verification | Replace seal per workflow and perform leak test and cycle verification | Leak test passes and cycle parameters acceptable after replacement | Leak test gauge; calibrated instruments | Per replacement | Maintenance; QA |
| OQ-135 | FRS-135: Carrier thermal mass mapping tests | Run mapping and BI/CI tests with each carrier type and record differences | Mapping and BI/CI acceptance met for each carrier type | Data loggers; BI certs | Per carrier type | Production; Validation Lead |
| OQ-136 | FRS-136: Adhesive/ink chemical analysis pre/post tests | Perform chemical analysis pre/post cycle and compare results to acceptance | No unacceptable migration or degradation observed | Chemical analysis lab certs | Per material | R&D; QA |
| OQ-137 | FRS-137: Nested configuration BI/CI and drying verification | Execute nested configuration tests with BIs at nested cores and drying checks | BI/CI negative at nested cores and drying acceptance met | BI certs; moisture balance | Per nested config | Production; Validation Lead |
| OQ-138 | FRS-138: Variable initial moisture tests and pre-conditioning verification | Run tests with variable initial moisture and verify drying acceptance or pre-conditioning effectiveness | Residual moisture within acceptance or pre-conditioning applied and effective | Moisture meter; moisture balance | Per test plan | Validation Lead; Production |
| OQ-139 | FRS-139: Containment and biosafety verification for biological loads | Execute specialized PQ with containment and verify biosafety controls and BI/CI results | Containment measures effective; BI/CI results acceptable; biosafety approvals recorded | Biosafety equipment certs | Per specialized PQ | R&D; QA; EHS |
| OQ-140 | FRS-140: Management review and lessons learned capture verification | Conduct lessons-learned review and verify VMP/URS update workflow | Management review recorded and VMP/URS updates captured with action tracking | Management review minutes | Once | QA; Validation Lead |

Deviation Handling during OQ:

Record any nonconformance on the deviation form. Assess impact on sterility, safety, and validated configuration. If safety or sterility risk exists, stop testing, initiate containment, and open CAPA. Rework, retest affected OQ items, and attach retest evidence and root-cause analysis.

OQ Signatures

Prepared by: Validation Lead — Date: ________

Reviewed by: Engineering Manager — Date: ________

Approved by: QA Manager — Date: ________

Performance Qualification PQ Protocol

Purpose:

Demonstrate the sterilization process meets the User Requirement Specifications (URS) under routine production conditions and that product release criteria, traceability, and operational workflows satisfy URS acceptance.

Prerequisites:

IQ and OQ complete and signed; calibrated instruments and data loggers on site; SOPs available; BI/CI lots and incubators ready; trained production staff; PQ load patterns and BI/CI placement plans defined; archive and export workflows validated.

Execution Notes:

Execute PQ as defined per URS mapping: run production representative loads, place BIs/CIs at worst-case positions, capture raw data and BI/CI incubation results, print and attach labels and certificates, and compile PQ report with traceability matrix. Record deviations immediately and route to CAPA if they affect sterility, safety, or release decisions. Retain raw data and chain of custody evidence.

PQ Protocol Table

| Test ID | Requirement URS mapping | Test Method | Acceptance Criterion | Instrument Calibration | Frequency | Responsible |
| --- | --- | --- | --- | --- | --- | --- |
| PQ-001 | URS-001 Terminal sterilization of all device types | Run 3 consecutive production loads for each device type with BI/CI at worst-case positions | All BIs negative; CIs indicate exposure; load records complete | Calibrated data loggers; BI/CI certs | Per PQ campaign per device type | Validation Lead; Production; QA |
| PQ-002 | URS-002 Support single/double wrap rigid containers peel pouches | Run production loads for each packaging format and inspect seals and drying | No seal failures; CIs pass; residual moisture acceptable | Moisture balance; CI/BI certs | Per packaging type PQ | Production; QA |
| PQ-003 | URS-003 Max gross load mass 50 kg | Run representative production loads at or near 50 kg and verify cycle performance | Cycle parameters within limits; BIs negative; no mechanical issues | Calibrated scale; data loggers | Per PQ campaign at max load | Production; Engineering |
| PQ-004 | URS-004 Chamber dimensions rack compatibility | Run production loads using specified racks and verify fit and steam access | Racks fit; no obstruction; BI/CI results acceptable | Visual; data loggers | Per rack type PQ | Production; Engineering |
| PQ-005 | URS-005 Pre-vacuum cycles 134°C/3 min and 132°C/4 min | Run 3 consecutive production loads on each validated recipe and record F₀ | BIs negative; F₀ meets URS lethality; cycle logs within OQ limits | Calibrated thermocouples; pressure transducers | Per recipe PQ | Validation Lead; Production |
| PQ-006 | URS-006 Bowie-Dick and dynamic air removal tests | Include BD or equivalent checks per PQ schedule and record results | BD passes per acceptance; no air removal anomalies in production loads | BD packs; vacuum gauge | Per PQ campaign or daily per SOP | Production; Validation Lead |
| PQ-007 | URS-007 User programmable cycles with role authorization | Verify no unauthorized recipe changes during PQ and audit trail entries for any edits | No unauthorized edits; any edits authorized and logged with reason | System audit log | During PQ campaign | IT; QA |
| PQ-008 | URS-008 Controller setpoint accuracy ±0.5°C | Review PQ cycle plateau data vs calibrated reference probes across loads | Plateau within ±0.5°C for production loads | Calibrated reference probes | Per PQ campaign | Validation Lead; QA |
| PQ-009 | URS-009 Temperature uniformity ±2.0°C | Verify loaded mapping points across representative production loads | All mapped points within ±2.0°C of setpoint | Calibrated data loggers | Per PQ campaign | Validation Lead |
| PQ-010 | URS-010 Temperature sensor redundancy | Monitor sensor discrepancy alarms during PQ and verify fallback behavior if triggered | No unresolved sensor discrepancies; fallback behavior logged and acceptable | Sensor calibration certs | During PQ campaign | Engineering; QA |
| PQ-011 | URS-011 Pressure and vacuum sensor accuracy ±2% FS | Review pressure/vacuum logs across PQ loads and compare to calibrated references | Readings within ±2% FS during production cycles | Calibrated pressure/vacuum gauges | Per PQ campaign | Engineering; QA |
| PQ-012 | URS-012 Vacuum capability ≤50 mbar leak rate ≤5 mbar/min | Verify vacuum performance and leak rate during production loaded cycles | Vacuum ≤50 mbar achieved and leak rate ≤5 mbar/min | Calibrated vacuum gauge | Per PQ campaign | Engineering |
| PQ-013 | URS-013 Steam quality dryness ≥0.95 NCG ≤3.5% | Collect steam samples during PQ and attach lab reports | Lab results meet dryness and NCG acceptance | Sampling kit; lab certs | Per PQ campaign or per steam source | Facilities; QA |
| PQ-014 | URS-014 Steam inlet pressure 3–6 bar | Monitor inlet pressure during PQ loads and verify within range | Inlet pressure maintained 3–6 bar during cycles | Calibrated pressure gauge | Continuous during PQ runs | Facilities |
| PQ-015 | URS-015 Data logging retention ≥10 years | Verify PQ raw data exported, checksum generated, and archived | Raw files archived with checksum and retrievable | Export tool; checksum utility | Per PQ campaign | IT; QA |
| PQ-016 | URS-016 Audit trail capture | Review audit trail for PQ activities including user IDs and approvals | Audit trail complete and immutable for PQ events | Audit DB printout | Per PQ campaign | IT; QA |
| PQ-017 | URS-017 Role based access control | Confirm only trained authorized users executed PQ runs and approvals | All PQ actions by authorized users; training records attached | Training matrix; system logs | Per PQ campaign | Training; QA |
| PQ-018 | URS-018 Electronic signature capability | Use e-signatures for PQ approvals where applicable and verify linkage | E-signatures present and linked to PQ records | System e-signature logs | Per PQ campaign where used | QA; IT |
| PQ-019 | URS-019 Secure export with checksum/digital signature | Export PQ cycle data and verify checksum/digital signature integrity | Exported PQ files verify against archive checksum | Checksum utility | Per PQ campaign | IT; QA |
| PQ-020 | URS-020 Alarm management and escalation | Review alarm logs during PQ and confirm operator responses per SOP | Alarms handled per SOP; escalation records present | Alarm logs; incident reports | During PQ campaign | Production; QA |
| PQ-021 | URS-021 Door interlock logging | Verify no door openings during unsafe conditions and review any attempts logged | No unsafe door openings; any attempts logged and investigated | Door sensor logs | During PQ campaign | Engineering; Safety |
| PQ-022 | URS-022 Emergency stop and fail-safe behavior | Confirm E-stop and power loss events during PQ are handled and data preserved | Any E-stop/power events resulted in safe state and data preserved | UPS logs; system logs | During PQ campaign | Engineering; IT |
| PQ-023 | URS-023 Firmware/software version control | Verify firmware/software version used during PQ matches approved V&V package | Firmware version matches approved V&V and recorded in PQ report | System printout; V&V docs | Per PQ campaign | IT; QA |
| PQ-024 | URS-024 Software validation and Part 11 controls | Confirm software controls used in PQ behaved per V&V and Part 11 where applicable | Software functions performed as validated; Part 11 controls enforced | Software test logs; audit trail | Per PQ campaign | IT; QA |
| PQ-025 | URS-025 Cycle reproducibility statistical control | Analyze PQ cycle data for reproducibility and SPC metrics | Key parameters within control limits; no out-of-control signals | SPC reports; data logger certs | Per PQ campaign | Validation Lead; QA |
| PQ-026 | URS-026 BI placement strategy and recording | Verify BIs placed per plan and results linked to load records for PQ loads | All BIs incubated and negative; placement recorded and traceable | BI lot certs; incubator logs | Per PQ load | Microbiology; Production |
| PQ-027 | URS-027 CI use and acceptance per pack/location | Inspect CIs from PQ loads and verify recorded results | CIs indicate exposure per acceptance and results logged | CI lot certs | Per PQ load | Production; QA |
| PQ-028 | URS-028 Residual moisture ≤3% | Measure residual moisture on PQ sample packs and record results | Residual moisture ≤3% or CI dry-pack pass for PQ samples | Moisture balance calibration cert | Per PQ load | Production; QA |
| PQ-029 | URS-029 Drying reproducibility | Review drying results across PQ loads for reproducibility | Drying results consistent and within defined SD limits | Moisture balance; SPC tool | Per PQ campaign | Validation Lead |
| PQ-030 | URS-030 Load configuration worst-case placement | Verify worst-case placements in PQ loads and mapping results | Worst-case positions meet T/P/vacuum acceptance and BIs negative | Data loggers; BI certs | Per PQ load pattern | Production; Validation Lead |
| PQ-031 | URS-031 Minimum PQ three consecutive production loads | Execute and document three consecutive successful production loads per cycle | Three consecutive loads pass all acceptance criteria | PQ workflow records | Per validated cycle | Validation Lead; QA |
| PQ-032 | URS-032 Requalification schedule adherence | Confirm any requalification triggers during PQ were handled per VMP | Requalification actions executed if triggered and documented | Change control records | As triggered | QA; Validation Lead |
| PQ-033 | URS-033 Spare parts availability during PQ | Verify no downtime during PQ due to missing critical spares | No PQ interruptions due to spare shortages | Spare inventory records | During PQ campaign | Procurement; Maintenance |
| PQ-034 | URS-034 PM performed without altering validated state | Review PM records during PQ period and confirm no unauthorized changes | PMs performed per schedule and no unauthorized parameter changes | CMMS records | During PQ campaign | Maintenance; QA |
| PQ-035 | URS-035 Calibration traceability for instruments used in PQ | Verify calibration certs for instruments used in PQ are current and include uncertainty | All instruments calibrated and certs attached with uncertainty fields | Calibration certificates | Per PQ campaign | Validation Lead; QA |
| PQ-036 | URS-036 Data logger mounting reproducibility | Confirm data logger mounting photos and positions match OQ documentation for PQ runs | Mounting photos attached and positions match OQ; mapping reproducible | Camera; data logger certs | Per PQ run | Validation Lead |
| PQ-037 | URS-037 Utility failure detection during PQ | Review utility alarm events during PQ and confirm appropriate response | Any utility failures handled per SOP and logged; no compromised loads released | Utility logs; incident reports | During PQ campaign | Facilities; QA |
| PQ-038 | URS-038 Network security posture during PQ | Confirm no critical security events affected PQ data integrity | No critical vulnerabilities exploited during PQ; logs reviewed | Vulnerability scan report | Per PQ campaign | IT; Security |
| PQ-039 | URS-039 Backup and restore verification for PQ data | Verify PQ raw data backups completed and restore tested for PQ dataset | PQ data backed up and restore verified; checksums match | Backup logs; checksum utility | Per PQ campaign | IT; QA |
| PQ-040 | URS-040 Archive retrieval for PQ records | Retrieve PQ package from archive and verify completeness | PQ package retrieved complete with attachments within SLA | Archive retrieval test | Per PQ campaign | QA; IT |
| PQ-041 | URS-041 Operator training competency during PQ | Confirm operators performing PQ have required training and competency records | All PQ operators trained and competency records attached | Training matrix | Per PQ campaign | Training; QA |
| PQ-042 | URS-042 SOP adherence during PQ | Verify SOP confirmations on HMI and checklist completion for PQ runs | SOP confirmations logged and checklists completed for each PQ run | HMI logs; checklists | Per PQ run | Production; QA |
| PQ-043 | URS-043 BI/CI lot control during PQ | Verify BI/CI lots used in PQ are in date and certificates attached | All BI/CI lots valid and certificates attached to PQ records | BI/CI certs | Per PQ load | Microbiology; QA |
| PQ-044 | URS-044 Positive and negative BI controls for PQ | Include positive/negative controls with PQ runs and verify incubation validity | Controls show expected results validating incubation | Incubator logs; control BI certs | Per PQ campaign | Microbiology |
| PQ-045 | URS-045 Load traceability to device lot numbers | Verify load records include device lot numbers, operator, cycle ID, and BI/CI results | Load records complete and scannable to device lots | Barcode scanner; system logs | Per PQ load | Production; QA |
| PQ-046 | URS-046 FAT/SAT evidence referenced in PQ report | Attach FAT/SAT deliverables to PQ report and confirm linkage | FAT/SAT evidence present and referenced in PQ report | Document repository printout | Per PQ campaign | Procurement; QA |
| PQ-047 | URS-047 Material certificates and surface finish verification | Inspect materials in PQ loads for compatibility evidence and surface finish integrity | No material incompatibility observed; certificates referenced | Material certs; visual inspection | Per PQ campaign | QA; Materials Eng |
| PQ-048 | URS-048 Cleaning verification prior to PQ | Confirm cleaning performed per SOP before PQ and cleaning logs attached | Cleaning logs present and visual cleanliness acceptable | Cleaning log; visual checklist | Before PQ runs | Production; QA |
| PQ-049 | URS-049 Ambient operating range during PQ | Verify ambient T/H during PQ loads and confirm no adverse effect on results | Ambient within URS range or mitigations documented; no adverse effect | Calibrated ambient sensors | During PQ campaign | Facilities; QA |
| PQ-050 | URS-050 Installation anchoring and clearance integrity during PQ | Confirm no installation issues observed during PQ operation | No movement or clearance violations during PQ | Visual inspection | During PQ campaign | Engineering |
| PQ-051 | URS-051 Condensate handling during PQ | Verify condensate routing and trap performance during PQ loads | No condensate backflow; traps functioning; no impact on drying | Visual; condensate logs | During PQ runs | Facilities |
| PQ-052 | URS-052 Audible and visual alarms escalation during PQ | Test alarm escalation during PQ and verify operator notification records | Alarms escalated per SOP and operator responses logged | Alarm logs; notification records | During PQ campaign | Production; QA |
| PQ-053 | URS-053 Partial requalification after minor repairs during PQ | If minor repairs occurred, verify partial requalification tests were executed and passed before resuming PQ | Partial requalification passed and documented before production resumed | Requalification records | As required | Validation Lead; QA |
| PQ-054 | URS-054 Raw data provenance for PQ | Verify raw PQ data files include metadata and checksum and are archived | Raw files include metadata and checksum and archive entry verified | Export tool; checksum utility | Per PQ campaign | IT; QA |
| PQ-055 | URS-055 Measurement uncertainty considered in PQ acceptance | Document how instrument uncertainty affected PQ acceptance decisions | Acceptance decisions include uncertainty propagation and justification | Calibration certs with uncertainty | During PQ analysis | Validation Lead; QA |
| PQ-056 | URS-056 Lumen worst-case PQ testing | Place BIs in lumens per lumen matrix during PQ loads and verify results | BIs negative at lumen tips for validated lumen dimensions | Lumen fixtures; BI certs | Per lumen matrix | Microbiology; Validation Lead |
| PQ-057 | URS-057 BD equivalents for production monitoring | Include BD or equivalent checks per PQ schedule and record results | BD passes per acceptance and production monitoring shows no air removal issues | BD packs; vacuum logs | Per PQ campaign | Production; Validation Lead |
| PQ-058 | URS-058 Change control impact during PQ | Verify any changes during PQ were processed via change control and requalification performed if required | All changes documented and requalification performed where triggered | Change control records | As occurred | QA |
| PQ-059 | URS-059 CAPA closure for PQ deviations | Review any PQ deviations and confirm CAPA actions closed with verification | Deviations resolved with CAPA and verification evidence attached | CAPA records | As required | QA |
| PQ-060 | URS-060 Final PQ validation report compilation | Compile PQ report with traceability matrix, raw data, BI/CI results, deviations, CAPA, and approvals | Final PQ report complete and signed with all attachments | Report generator output | After PQ campaign | Validation Lead; QA |
| PQ-061 | URS-061 Material compatibility evidence in PQ | Verify materials in PQ loads show no unacceptable degradation post-cycle | No unacceptable degradation; test evidence attached | Material test reports | Per material | R&D; QA |
| PQ-062 | URS-062 Packaging matrix PQ verification | Run PQ loads per packaging matrix and verify seal integrity and CI/BI outcomes | Packaging integrity maintained and CI/BI pass for PQ loads | Peel test results; CI/BI certs | Per packaging matrix | Packaging Eng; QA |
| PQ-063 | URS-063 Lumen and geometry PQ verification | Execute PQ loads covering lumen/geometry matrix and verify BI results | BI negative for validated geometries and nested configs | BI certs; mapping logs | Per matrix | Microbiology; Validation Lead |
| PQ-064 | URS-064 Assembled and disassembled device PQ verification | Run PQ loads for both assembly states and perform functional checks post-cycle | Both states meet BI/CI and functional acceptance criteria | Functional test records; BI certs | Per device | Production; QA |
| PQ-065 | URS-065 Stability sampling during PQ | Collect baseline stability samples during PQ and record storage conditions | Stability samples collected and baseline tests meet acceptance | Stability sample logs | Per PQ campaign | Regulatory Affairs; QA |
| PQ-066 | URS-066 Feedwater and condensate quality during PQ | Review water/condensate lab results during PQ and confirm within limits | Water/condensate results within URS limits during PQ | Lab certificates | Per PQ campaign | Facilities; QA |
| PQ-067 | URS-067 Altitude correction verification in PQ | Confirm PQ cycles at site altitude meet BI acceptance with applied corrections | BI negative and adjusted parameters documented in PQ report | Controller logs; BI certs | Per PQ campaign | Validation Lead; Engineering |
| PQ-068 | URS-068 HVAC environmental linkage for storage PQ | Verify storage environmental data linked to PQ lots and no excursions affecting sterility | Environmental data within limits and linked to lot records | Environmental logs | During PQ campaign | Facilities; QA |
| PQ-069 | URS-069 Environmental compliance sampling during PQ | Attach condensate/exhaust sample reports to PQ and confirm compliance where applicable | Sample results documented and compared to permit limits | Lab certs | Per PQ campaign | Facilities; EHS |
| PQ-070 | URS-070 Patch management during PQ | Confirm no unapproved patches applied during PQ and any patches tested per policy | No production patches applied during PQ or approved and tested per change control | Patch logs; change control records | During PQ campaign | IT; QA |
| PQ-071 | URS-071 Time synchronization integrity during PQ | Verify timestamps on PQ records match NTP and show no unacceptable drift | Timestamps synchronized and drift within tolerance for PQ records | NTP logs | Per PQ campaign | IT; QA |
| PQ-072 | URS-072 Data encryption and secure transfer during PQ | Confirm PQ data transfers were encrypted and logs show secure transfers | Secure transfer logs present and encryption verified | Security logs | Per PQ campaign | IT; Security |
| PQ-073 | URS-073 MES/LIMS integration for PQ release | Verify PQ release status transmitted to MES/LIMS and reconciliation successful | MES/LIMS receives release only after BI/CI pass and reconciliation matches | API logs; reconciliation report | Per PQ release | IT; QA |
| PQ-074 | URS-074 Disaster recovery verification for PQ data | Confirm PQ dataset recoverable within RTO/RPO in restore test | PQ data restored within RTO/RPO and integrity verified | Backup/restore logs | Per PQ campaign | IT |
| PQ-075 | URS-075 Lot release rules enforcement during PQ | Verify automated hold/release logic prevented release until BI/CI and approvals complete | No loads released without BI/CI pass and approvals; holds functioned | Release logs; MES handoff | Per PQ release | QA; IT |
| PQ-076 | URS-076 Statistical sampling plan application in PQ | Apply sampling plan to PQ batch releases and document acceptance decisions | Sampling results recorded and acceptance decisions per plan | Sampling records; statistical tool | Per batch release | QA; Biostatistics |
| PQ-077 | URS-077 Indicator selection enforcement during PQ | Confirm indicators used in PQ match selection rules and lot controls | Indicators conform to master data and lot certificates attached | Indicator master data; certs | Per PQ load | Microbiology; QA |
| PQ-078 | URS-078 Environmental monitoring trend review during PQ | Review environmental trends for storage during PQ period and document findings | Trend review shows no adverse trends affecting PQ lots | Trend reports | Per PQ campaign | QA; Facilities |
| PQ-079 | URS-079 Labeling and traceability verification for PQ loads | Verify printed labels and electronic traceability fields for PQ loads match records | Labels accurate and load records link to printed labels and device lots | Label print logs; barcode scans | Per PQ load | Production; QA |
| PQ-080 | URS-080 Critical spare availability during PQ | Confirm critical spares were available and no PQ interruptions occurred due to spares | No PQ downtime due to missing spares; inventory records verified | Spare inventory | During PQ campaign | Procurement; Maintenance |
| PQ-081 | URS-081 Maintenance trigger compliance during PQ | Verify any maintenance during PQ triggered required requalification and was documented | Maintenance triggers processed and requalification performed if required | Maintenance records; change control | As occurred | Maintenance; QA |
| PQ-082 | URS-082 Consumable qualification enforcement during PQ | Confirm consumables used in PQ are qualified and certificates attached | All consumables qualified and certificates attached to PQ records | Consumable certs | Per PQ load | Procurement; QA |
| PQ-083 | URS-083 Decommissioning records not applicable during PQ | N/A for PQ unless decommissioning occurs | N/A | N/A | N/A | QA |
| PQ-084 | URS-084 Regulatory reporting triggers during PQ | Verify system flagged any reportable events during PQ and notifications executed | Any reportable events identified and notifications executed per timelines | Event logs; notification records | As occurred | Regulatory Affairs; QA |
| PQ-085 | URS-085 Emergency procedures readiness during PQ | Confirm emergency SOPs and drill records available and staff aware during PQ | Emergency procedures accessible and staff trained; drills documented | Drill logs; SOPs | Per PQ campaign | QA; EHS |
| PQ-086 | URS-086 Audit pack readiness for PQ | Generate PQ audit pack and verify completeness for inspection | Audit pack contains required PQ evidence and retrievable within SLA | Audit pack generator output | Per PQ campaign | QA |
| PQ-087 | URS-087 Local regulatory compliance during PQ | Verify permits valid and compliance evidence attached for PQ operations | Permits valid during PQ and compliance evidence attached | Permit repository | Per PQ campaign | Regulatory Affairs |
| PQ-088 | URS-088 Consumable/spare substitution control during PQ | Confirm any substitutions used in PQ were approved via change control and verified | Substitutions approved and verification tests passed | Change control records | As occurred | Procurement; QA |
| PQ-089 | URS-089 Training requalification frequency compliance during PQ | Verify personnel performing PQ have up-to-date requalification records | All PQ personnel meet requalification requirements | Training matrix | Per PQ campaign | Training; QA |
| PQ-090 | URS-090 Trending and management review evidence from PQ | Include PQ results in trend reports and management review actions | PQ results included in trend analysis and management review actions recorded | Trend reports; review minutes | Per PQ campaign | QA |
| PQ-091 | URS-091 Emergency cycle validation evidence during PQ | If emergency cycles used, verify limited validation and approvals documented | Emergency cycle used only with documented limited validation and approvals | Emergency cycle records | As occurred | Validation Lead; QA |
| PQ-092 | URS-092 Third-party service provider records during PQ | Verify any third-party work during PQ performed by qualified providers and records attached | Service reports attached and provider qualification on file | Service reports | As occurred | Maintenance; Procurement |
| PQ-093 | URS-093 Accessory validation included in PQ loads | Verify accessories included in PQ loads and BI/CI results acceptable | Accessories do not adversely affect BI/CI or mapping; results acceptable | BI certs; mapping logs | Per accessory | Production; Validation Lead |
| PQ-094 | URS-094 Nonconforming product handling during PQ | Confirm any nonconforming PQ product handled per disposition workflow | Nonconforming items quarantined and disposition documented | Nonconforming module logs | As occurred | QA |
| PQ-095 | URS-095 Trap/separator maintenance verification during PQ | Verify trap/separator maintenance performed and steam quality acceptable for PQ loads | Post-maintenance steam quality acceptable and records attached | Sampling kit; lab certs | As scheduled | Facilities; Maintenance |
| PQ-096 | URS-096 MSA evidence for instruments used in PQ | Include MSA/Gage R&R evidence for instruments used in PQ analysis | Instruments used have acceptable MSA results attached to PQ report | MSA reports | Per PQ campaign | Validation Lead; QA |
| PQ-097 | URS-097 VMP review evidence related to PQ outcomes | Document any VMP updates resulting from PQ lessons learned | VMP updated if required and management review minutes attached | Management review record | Post PQ campaign | QA |
| PQ-098 | URS-098 Supplier deliverable evidence in PQ report | Attach supplier FAT/SAT and spares evidence used to support PQ | Supplier deliverables present and referenced in PQ report | Supplier docs | Per PQ campaign | Procurement; QA |
| PQ-099 | URS-099 Seasonal ambient performance evidence in PQ | Include PQ evidence for seasonal extremes or documented mitigations | PQ loads meet acceptance under extremes or mitigations effective | Environmental test records | Per site needs | Facilities; Validation Lead |
| PQ-100 | URS-100 Measurement uncertainty documentation in PQ decisions | Document how measurement uncertainty influenced PQ acceptance decisions | Uncertainty propagation documented and acceptance justified | Calibration certs with uncertainty | During PQ analysis | Validation Lead; QA |
| PQ-101 | URS-101 Pre-cleaning and bioburden verification for PQ loads | Verify pre-cleaning performed and incoming bioburden within limits for PQ loads | Pre-cleaning logs attached and bioburden results within limits | Microbiology lab certs | Per PQ load | Production; Microbiology |
| PQ-102 | URS-102 Approved detergents usage verification during PQ | Confirm only approved detergents used prior to PQ and compatibility evidence attached | Only approved detergents used and compatibility evidence present | Cleaning logs; compatibility certs | Per PQ campaign | Production; QA |
| PQ-103 | URS-103 Steam source validation evidence in PQ | If multiple steam sources used, include PQ evidence for each source used | PQ loads meet acceptance for each steam source used | Steam sample lab certs | Per steam source used | Facilities; Validation Lead |
| PQ-104 | URS-104 UPS and emergency power behavior during PQ | Verify no data loss or unsafe states during any power events in PQ | No data loss and safe state maintained during any power events | UPS logs; system logs | During PQ campaign | IT; Facilities |
| PQ-105 | URS-105 EMC and grounding evidence during PQ | Confirm no EMI anomalies affected PQ runs and grounding maintained | No EMI anomalies observed and grounding checks passed | EMC/grounding check logs | Per PQ campaign | Engineering; IT |
| PQ-106 | URS-106 Noise and vibration evidence during PQ | Verify noise/vibration remained within limits during PQ | Noise/vibration within limits and no operator safety issues | Sound/vibration meter certs | During PQ campaign | Facilities; EHS |
| PQ-107 | URS-107 Seasonal ambient mitigation evidence during PQ | Include PQ evidence showing acceptable performance or mitigations for seasonal extremes | PQ evidence shows acceptable performance or mitigations effective | Environmental test records | Per site needs | Facilities; Validation Lead |
| PQ-108 | URS-108 MSA repeatability/reproducibility evidence in PQ | Include MSA/Gage R&R evidence for instruments used in PQ | MSA results acceptable and attached to PQ report | MSA reports | Per PQ campaign | Validation Lead; QA |
| PQ-109 | URS-109 Sensor drift monitoring during PQ | Review sensor drift logs during PQ and confirm recalibration actions if triggered | No unaddressed drift; recalibration actions recorded if thresholds exceeded | Calibration certs; drift logs | During PQ campaign | Engineering; QA |
| PQ-110 | URS-110 Firmware rollback evidence if used during PQ | If rollback occurred, document controlled rollback and data integrity verification | Rollback documented and data integrity verified; approvals attached | Firmware repo logs | As occurred | IT; QA |
| PQ-111 | URS-111 Remote vendor access records during PQ | Verify any vendor remote sessions during PQ were authorized and logged | Remote sessions authorized and session logs attached | Remote access logs | As occurred | IT; QA |
| PQ-112 | URS-112 Mobile device access audit during PQ | Confirm any mobile access during PQ was authorized and logged | Mobile actions authenticated and audit logs present | Mobile device logs | During PQ campaign | IT; QA |
| PQ-113 | URS-113 MES/LIMS release interlock evidence during PQ | Verify automated release to MES/LIMS occurred only after BI/CI pass and approvals | MES/LIMS release records match PQ approvals and BI/CI results | MES/LIMS logs; reconciliation report | Per PQ release | QA; IT |
| PQ-114 | URS-114 Sample retention verification for PQ samples | Confirm retained PQ samples stored per retention policy and retrievable | Retained samples logged with location and retrieval tested | Retention system logs | Per PQ campaign | QA |
| PQ-115 | URS-115 Post-maintenance verification evidence during PQ | If maintenance occurred, confirm post-maintenance verification passed before PQ loads released | Post-maintenance verification passed and sign-off attached | Verification scripts; records | As occurred | Maintenance; Validation Lead |
| PQ-116 | URS-116 Obsolescence and alternative part evidence during PQ | If alternative parts used in PQ, confirm change control and verification tests passed | Alternative parts approved and verification evidence attached | Change control records | As occurred | Procurement; QA |
| PQ-117 | URS-117 Orientation and stacking PQ verification | Verify PQ loads executed validated orientations/stacking and BI/CI results acceptable | BI/CI negative and mapping acceptable for validated orientations | BI certs; mapping logs | Per orientation pattern | Production; Validation Lead |
| PQ-118 | URS-118 Mixed-load PQ verification | Execute PQ mixed loads and verify BI/CI and mapping results across device types | BI/CI negative and mapping acceptable for mixed loads | BI certs; data loggers | Per mixed-load pattern | Production; Validation Lead |
| PQ-119 | URS-119 Implant handling PQ evidence | For implant loads, confirm enhanced handling, documentation, and regulatory sign-off | Implant PQ documentation complete and regulatory checklist signed | Implant workflow records | Per implant load | Regulatory Affairs; QA |
| PQ-120 | URS-120 Endotoxin control evidence during PQ | If endotoxin testing required, include results for PQ lots and confirm acceptance | Endotoxin results within acceptance where required and attached | Endotoxin lab certs | Per requirement | QA; Microbiology |
| PQ-121 | URS-121 Post-relocation PQ verification | After relocation, confirm post-move IQ/OQ passed and PQ evidence attached | Post-move IQ/OQ passed and PQ evidence attached before production | Post-move IQ/OQ records | After relocation | Validation Lead; Engineering |
| PQ-122 | URS-122 Multi-site comparability evidence in PQ | For multi-site deployments, include cross-site comparability evidence in PQ report | Cross-site comparability demonstrated or deviations justified | Cross-site reports | Per multi-site PQ | Corporate Validation; QA |
| PQ-123 | URS-123 Supplier audit evidence used in PQ | Confirm supplier deliverables used in PQ are from audited suppliers and records attached | Supplier audit evidence present and deliverables accepted | Supplier audit reports | Per PQ campaign | Procurement; QA |
| PQ-124 | URS-124 Accessory cleaning PQ inclusion evidence | Verify accessories included in PQ loads passed cleaning validation and BI/CI acceptance | Accessories passed cleaning validation and PQ BI/CI acceptable | Cleaning validation records; BI certs | Per accessory | Production; Validation Lead |
| PQ-125 | URS-125 Raw data provenance chain of custody for PQ | Verify chain of custody and provenance metadata for PQ raw files | Raw files include provenance metadata and chain of custody entries | Export tool; chain of custody logs | Per PQ campaign | IT; QA |
| PQ-126 | URS-126 Redaction workflow evidence for external PQ sharing | If PQ records shared externally, verify redaction preserved validation evidence and approval logs | Redacted records retain validation evidence and redaction approvals logged | Redaction logs | As occurred | QA; Regulatory Affairs |
| PQ-127 | URS-127 Water hardness mitigation evidence during PQ | If water hardness outside limits, verify mitigation controls applied and PQ acceptance | Mitigations applied and PQ acceptance met with documented evidence | Hardness test certs | As required | Facilities; QA |
| PQ-128 | URS-128 Trap/separator maintenance verification during PQ | Confirm trap/separator maintenance performed and steam quality acceptable for PQ loads | Post-maintenance steam quality acceptable and records attached | Sampling kit; lab certs | As scheduled | Facilities; Maintenance |
| PQ-129 | URS-129 Humidity variation drying evidence during PQ | Include PQ evidence showing drying acceptance across humidity range or mitigations | Residual moisture acceptable across humidity range or mitigations documented | Humidity sensors; moisture balance | Per PQ campaign | Validation Lead; Facilities |
| PQ-130 | URS-130 Embedded electronics functional verification in PQ | For loads with electronics, include pre/post functional test results in PQ report | Electronics function per vendor criteria post-PQ | Functional test records | Per device | R&D; QA |
| PQ-131 | URS-131 CI class interpretation evidence in PQ | Document CI class used and interpretation per SOP for PQ loads | CI interpretation recorded per SOP and linked to PQ report | CI certs; SOP | Per PQ load | Microbiology; QA |
| PQ-132 | URS-132 Alternative cycle PQ evidence | If alternative cycles used in production, include PQ evidence and material compatibility results | Alternative cycle PQ passes BI/CI and material compatibility acceptance | BI certs; material test reports | Per alternative cycle | Validation Lead; R&D |
| PQ-133 | URS-133 Timestamp integrity evidence in PQ | Verify timestamps on PQ records correct and synchronized to NTP | Timestamps synchronized and verification logs attached | NTP logs | Per PQ campaign | IT; QA |
| PQ-134 | URS-134 Seal replacement PQ verification | If seals replaced prior to PQ, include leak test and PQ evidence post-replacement | Leak test passed and PQ loads acceptable after replacement | Leak test gauge; BI certs | As occurred | Maintenance; QA |
| PQ-135 | URS-135 Carrier thermal mass PQ evidence | Include PQ results for each carrier type used in production | Carrier types validated with acceptable mapping and BI/CI results | Data loggers; BI certs | Per carrier type | Production; Validation Lead |
| PQ-136 | URS-136 Adhesive/ink migration evidence in PQ | If adhesives/inks present, include pre/post chemical analysis results in PQ report | No unacceptable migration or degradation observed | Chemical analysis lab certs | Per PQ campaign | R&D; QA |
| PQ-137 | URS-137 Nested packaging PQ evidence | Include PQ results for nested configurations with BI/CI at nested cores | Nested cores BI/CI negative and drying acceptable | BI certs; moisture balance | Per nested config | Production; Validation Lead |
| PQ-138 | URS-138 Variable initial moisture PQ evidence | If variable initial moisture present, include PQ evidence showing acceptance or pre-conditioning effectiveness | Residual moisture acceptable or pre-conditioning applied and effective | Moisture meter; moisture balance | Per PQ campaign | Validation Lead; Production |
| PQ-139 | URS-139 Biological materials PQ containment evidence | For biological loads, include containment, biosafety approvals, and PQ results | Containment effective; BI/CI results acceptable; approvals attached | Biosafety records | Per specialized PQ | R&D; QA; EHS |
| PQ-140 | URS-140 Continuous improvement capture from PQ | Document lessons learned from PQ and VMP/URS updates resulting from PQ | Lessons learned captured and VMP/URS updated with action closure evidence | Management review minutes | Post PQ campaign | QA; Validation Lead |

Deviation Handling during PQ:

Record any nonconformance on the deviation form. Assess impact on sterility, safety, and product release. If safety or sterility risk exists, stop production, initiate containment, and open CAPA. Rework, retest affected PQ items, and attach retest evidence and root-cause analysis.

PQ Signatures:

Prepared by: Validation Lead — Date: ________

Reviewed by: Production Manager — Date: ________

Approved by: QA Manager — Date: ________

IQ / OQ / PQ Traceability Matrix (URS → FRS → TRS → IQ / OQ / PQ)

| URS ID | FRS ID | TRS ID | IQ Reference | OQ Reference | PQ Reference |
| --- | --- | --- | --- | --- | --- |
| URS-001 | FRS-001 | TRS-001 | IQ-01 | OQ-001 | PQ-001 |
| URS-002 | FRS-002 | TRS-002 | IQ-02 | OQ-002 | PQ-002 |
| URS-003 | FRS-003 | TRS-003 | IQ-03 | OQ-003 | PQ-003 |
| URS-004 | FRS-004 | TRS-004 | IQ-03 | OQ-004 | PQ-004 |
| URS-005 | FRS-005 | TRS-005 | IQ-01 | OQ-005 | PQ-005 |
| URS-006 | FRS-006 | TRS-006 | IQ-16 | OQ-006 | PQ-006 |
| URS-007 | FRS-007 | TRS-007 | IQ-09 | OQ-007 | PQ-007 |
| URS-008 | FRS-008 | TRS-008 | IQ-08 | OQ-008 | PQ-008 |
| URS-009 | FRS-009 | TRS-009 | IQ-08 | OQ-009 | PQ-009 |
| URS-010 | FRS-010 | TRS-010 | IQ-08 | OQ-010 | PQ-010 |
| URS-011 | FRS-011 | TRS-011 | IQ-08 | OQ-011 | PQ-011 |
| URS-012 | FRS-012 | TRS-012 | IQ-04 | OQ-012 | PQ-012 |
| URS-013 | FRS-013 | TRS-013 | IQ-05 | OQ-013 | PQ-013 |
| URS-014 | FRS-014 | TRS-014 | IQ-04 | OQ-014 | PQ-014 |
| URS-015 | FRS-015 | TRS-015 | IQ-09 | OQ-015 | PQ-015 |
| URS-016 | FRS-016 | TRS-016 | IQ-09 | OQ-016 | PQ-016 |
| URS-017 | FRS-017 | TRS-017 | IQ-12 | OQ-017 | PQ-017 |
| URS-018 | FRS-018 | TRS-018 | IQ-09 | OQ-018 | PQ-018 |
| URS-019 | FRS-019 | TRS-019 | IQ-09 | OQ-019 | PQ-019 |
| URS-020 | FRS-020 | TRS-020 | IQ-14 | OQ-020 | PQ-020 |
| URS-021 | FRS-021 | TRS-021 | IQ-07 | OQ-021 | PQ-021 |
| URS-022 | FRS-022 | TRS-022 | IQ-07 | OQ-022 | PQ-022 |
| URS-023 | FRS-023 | TRS-023 | IQ-01 | OQ-023 | PQ-023 |
| URS-024 | FRS-024 | TRS-024 | IQ-02 | OQ-024 | PQ-024 |
| URS-025 | FRS-025 | TRS-025 | IQ-01 | OQ-025 | PQ-025 |
| URS-026 | FRS-026 | TRS-026 | IQ-17 | OQ-026 | PQ-026 |
| URS-027 | FRS-027 | TRS-027 | IQ-17 | OQ-027 | PQ-027 |
| URS-028 | FRS-028 | TRS-028 | IQ-28 (site IQ task) | OQ-028 | PQ-028 |
| URS-029 | FRS-029 | TRS-029 | IQ-28 (site IQ task) | OQ-029 | PQ-029 |
| URS-030 | FRS-030 | TRS-030 | IQ-17 | OQ-030 | PQ-030 |
| URS-031 | FRS-031 | TRS-031 | IQ-13 (readiness) | OQ-031 | PQ-031 |
| URS-032 | FRS-032 | TRS-032 | IQ-02 | OQ-032 | PQ-032 |
| URS-033 | FRS-033 | TRS-033 | IQ-11 | OQ-033 | PQ-033 |
| URS-034 | FRS-034 | TRS-034 | IQ-11 | OQ-034 | PQ-034 |
| URS-035 | FRS-035 | TRS-035 | IQ-10 | OQ-035 | PQ-035 |
| URS-036 | FRS-036 | TRS-036 | IQ-17 | OQ-036 | PQ-036 |
| URS-037 | FRS-037 | TRS-037 | IQ-04 | OQ-037 | PQ-037 |
| URS-038 | FRS-038 | TRS-038 | IQ-09 | OQ-038 | PQ-038 |
| URS-039 | FRS-039 | TRS-039 | IQ-15 | OQ-039 | PQ-039 |
| URS-040 | FRS-040 | TRS-040 | IQ-15 | OQ-040 | PQ-040 |
| URS-041 | FRS-041 | TRS-041 | IQ-12 | OQ-041 | PQ-041 |
| URS-042 | FRS-042 | TRS-042 | IQ-12 | OQ-042 | PQ-042 |
| URS-043 | FRS-043 | TRS-043 | IQ-11 | OQ-043 | PQ-043 |
| URS-044 | FRS-044 | TRS-044 | IQ-16 | OQ-044 | PQ-044 |
| URS-045 | FRS-045 | TRS-045 | IQ-18 | OQ-045 | PQ-045 |
| URS-046 | FRS-046 | TRS-046 | IQ-02 | OQ-046 | PQ-046 |
| URS-047 | FRS-047 | TRS-047 | IQ-02 | OQ-047 | PQ-047 |
| URS-048 | FRS-048 | TRS-048 | IQ-02 | OQ-048 | PQ-048 |
| URS-049 | FRS-049 | TRS-049 | IQ-19 | OQ-049 | PQ-049 |
| URS-050 | FRS-050 | TRS-050 | IQ-03 | OQ-050 | PQ-050 |
| URS-051 | FRS-051 | TRS-051 | IQ-06 | OQ-051 | PQ-051 |
| URS-052 | FRS-052 | TRS-052 | IQ-14 | OQ-052 | PQ-052 |
| URS-053 | FRS-053 | TRS-053 | IQ-02 | OQ-053 | PQ-053 |
| URS-054 | FRS-054 | TRS-054 | IQ-15 | OQ-054 | PQ-054 |
| URS-055 | FRS-055 | TRS-055 | IQ-10 | OQ-055 | PQ-055 |
| URS-056 | FRS-056 | TRS-056 | IQ-16 | OQ-056 | PQ-056 |
| URS-057 | FRS-057 | TRS-057 | IQ-16 | OQ-057 | PQ-057 |
| URS-058 | FRS-058 | TRS-058 | IQ-02 | OQ-058 | PQ-058 |
| URS-059 | FRS-059 | TRS-059 | IQ-02 | OQ-059 | PQ-059 |
| URS-060 | FRS-060 | TRS-060 | IQ-02 | OQ-060 | PQ-060 |
| URS-061 | FRS-061 | TRS-061 | IQ-02 | OQ-061 | PQ-061 |
| URS-062 | FRS-062 | TRS-062 | IQ-02 | OQ-062 | PQ-062 |
| URS-063 | FRS-063 | TRS-063 | IQ-17 | OQ-063 | PQ-063 |
| URS-064 | FRS-064 | TRS-064 | IQ-17 | OQ-064 | PQ-064 |
| URS-065 | FRS-065 | TRS-065 | IQ-17 | OQ-065 | PQ-065 |
| URS-066 | FRS-066 | TRS-066 | IQ-04 | OQ-066 | PQ-066 |
| URS-067 | FRS-067 | TRS-067 | IQ-01 | OQ-067 | PQ-067 |
| URS-068 | FRS-068 | TRS-068 | IQ-19 | OQ-068 | PQ-068 |
| URS-069 | FRS-069 | TRS-069 | IQ-06 | OQ-069 | PQ-069 |
| URS-070 | FRS-070 | TRS-070 | IQ-09 | OQ-070 | PQ-070 |
| URS-071 | FRS-071 | TRS-071 | IQ-09 | OQ-071 | PQ-071 |
| URS-072 | FRS-072 | TRS-072 | IQ-09 | OQ-072 | PQ-072 |
| URS-073 | FRS-073 | TRS-073 | IQ-09 | OQ-073 | PQ-073 |
| URS-074 | FRS-074 | TRS-074 | IQ-15 | OQ-074 | PQ-074 |
| URS-075 | FRS-075 | TRS-075 | IQ-09 | OQ-075 | PQ-075 |
| URS-076 | FRS-076 | TRS-076 | IQ-02 | OQ-076 | PQ-076 |
| URS-077 | FRS-077 | TRS-077 | IQ-11 | OQ-077 | PQ-077 |
| URS-078 | FRS-078 | TRS-078 | IQ-19 | OQ-078 | PQ-078 |
| URS-079 | FRS-079 | TRS-079 | IQ-18 | OQ-079 | PQ-079 |
| URS-080 | FRS-080 | TRS-080 | IQ-11 | OQ-080 | PQ-080 |
| URS-081 | FRS-081 | TRS-081 | IQ-02 | OQ-081 | PQ-081 |
| URS-082 | FRS-082 | TRS-082 | IQ-11 | OQ-082 | PQ-082 |
| URS-083 | FRS-083 | TRS-083 | IQ-02 | OQ-083 | PQ-083 |
| URS-084 | FRS-084 | TRS-084 | IQ-02 | OQ-084 | PQ-084 |
| URS-085 | FRS-085 | TRS-085 | IQ-02 | OQ-085 | PQ-085 |
| URS-086 | FRS-086 | TRS-086 | IQ-02 | OQ-086 | PQ-086 |
| URS-087 | FRS-087 | TRS-087 | IQ-02 | OQ-087 | PQ-087 |
| URS-088 | FRS-088 | TRS-088 | IQ-02 | OQ-088 | PQ-088 |
| URS-089 | FRS-089 | TRS-089 | IQ-12 | OQ-089 | PQ-089 |
| URS-090 | FRS-090 | TRS-090 | IQ-02 | OQ-090 | PQ-090 |
| URS-091 | FRS-091 | TRS-091 | IQ-02 | OQ-091 | PQ-091 |
| URS-092 | FRS-092 | TRS-092 | IQ-02 | OQ-092 | PQ-092 |
| URS-093 | FRS-093 | TRS-093 | IQ-02 | OQ-093 | PQ-093 |
| URS-094 | FRS-094 | TRS-094 | IQ-02 | OQ-094 | PQ-094 |
| URS-095 | FRS-095 | TRS-095 | IQ-06 | OQ-095 | PQ-095 |
| URS-096 | FRS-096 | TRS-096 | IQ-10 | OQ-096 | PQ-096 |
| URS-097 | FRS-097 | TRS-097 | IQ-02 | OQ-097 | PQ-097 |
| URS-098 | FRS-098 | TRS-098 | IQ-02 | OQ-098 | PQ-098 |
| URS-099 | FRS-099 | TRS-099 | IQ-02 | OQ-099 | PQ-099 |
| URS-100 | FRS-100 | TRS-100 | IQ-10 | OQ-100 | PQ-100 |
| URS-101 | FRS-101 | TRS-101 | IQ-02 | OQ-101 | PQ-101 |
| URS-102 | FRS-102 | TRS-102 | IQ-02 | OQ-102 | PQ-102 |
| URS-103 | FRS-103 | TRS-103 | IQ-05 | OQ-103 | PQ-103 |
| URS-104 | FRS-104 | TRS-104 | IQ-07 | OQ-104 | PQ-104 |
| URS-105 | FRS-105 | TRS-105 | IQ-03 | OQ-105 | PQ-105 |
| URS-106 | FRS-106 | TRS-106 | IQ-19 | OQ-106 | PQ-106 |
| URS-107 | FRS-107 | TRS-107 | IQ-19 | OQ-107 | PQ-107 |
| URS-108 | FRS-108 | TRS-108 | IQ-10 | OQ-108 | PQ-108 |
| URS-109 | FRS-109 | TRS-109 | IQ-08 | OQ-109 | PQ-109 |
| URS-110 | FRS-110 | TRS-110 | IQ-09 | OQ-110 | PQ-110 |
| URS-111 | FRS-111 | TRS-111 | IQ-09 | OQ-111 | PQ-111 |
| URS-112 | FRS-112 | TRS-112 | IQ-09 | OQ-112 | PQ-112 |
| URS-113 | FRS-113 | TRS-113 | IQ-09 | OQ-113 | PQ-113 |
| URS-114 | FRS-114 | TRS-114 | IQ-02 | OQ-114 | PQ-114 |
| URS-115 | FRS-115 | TRS-115 | IQ-02 | OQ-115 | PQ-115 |
| URS-116 | FRS-116 | TRS-116 | IQ-02 | OQ-116 | PQ-116 |
| URS-117 | FRS-117 | TRS-117 | IQ-17 | OQ-117 | PQ-117 |
| URS-118 | FRS-118 | TRS-118 | IQ-17 | OQ-118 | PQ-118 |
| URS-119 | FRS-119 | TRS-119 | IQ-12 | OQ-119 | PQ-119 |
| URS-120 | FRS-120 | TRS-120 | IQ-02 | OQ-120 | PQ-120 |
| URS-121 | FRS-121 | TRS-121 | IQ-03 | OQ-121 | PQ-121 |
| URS-122 | FRS-122 | TRS-122 | IQ-02 | OQ-122 | PQ-122 |
| URS-123 | FRS-123 | TRS-123 | IQ-02 | OQ-123 | PQ-123 |
| URS-124 | FRS-124 | TRS-124 | IQ-17 | OQ-124 | PQ-124 |
| URS-125 | FRS-125 | TRS-125 | IQ-15 | OQ-125 | PQ-125 |
| URS-126 | FRS-126 | TRS-126 | IQ-02 | OQ-126 | PQ-126 |
| URS-127 | FRS-127 | TRS-127 | IQ-04 | OQ-127 | PQ-127 |
| URS-128 | FRS-128 | TRS-128 | IQ-06 | OQ-128 | PQ-128 |
| URS-129 | FRS-129 | TRS-129 | IQ-19 | OQ-129 | PQ-129 |
| URS-130 | FRS-130 | TRS-130 | IQ-17 | OQ-130 | PQ-130 |
| URS-131 | FRS-131 | TRS-131 | IQ-17 | OQ-131 | PQ-131 |
| URS-132 | FRS-132 | TRS-132 | IQ-02 | OQ-132 | PQ-132 |
| URS-133 | FRS-133 | TRS-133 | IQ-09 | OQ-133 | PQ-133 |
| URS-134 | FRS-134 | TRS-134 | IQ-03 | OQ-134 | PQ-134 |
| URS-135 | FRS-135 | TRS-135 | IQ-17 | OQ-135 | PQ-135 |
| URS-136 | FRS-136 | TRS-136 | IQ-02 | OQ-136 | PQ-136 |
| URS-137 | FRS-137 | TRS-137 | IQ-17 | OQ-137 | PQ-137 |
| URS-138 | FRS-138 | TRS-138 | IQ-17 | OQ-138 | PQ-138 |
| URS-139 | FRS-139 | TRS-139 | IQ-17 | OQ-139 | PQ-139 |
| URS-140 | FRS-140 | TRS-140 | IQ-13 | OQ-140 | PQ-140 |